## Slide 1: Larson Davis Intrinsic Safety Training Agenda
Larson Davis Intrinsic Safety Products
Schedule Drawings / I.S. Part Numbering
Other Resources

## Slide 2: Larson Davis Intrinsic Safety Products
Currently Larson Davis has two families of products that have hazardous location (explosive atmosphere) safety ratings:
Spartan 730IS Noise Dosimeter (ETL [UL,CSA], ATEX, IECEx)
Spark noise dosimeters:
  703+, 705+, and 706RC; (UL, CSA, FM, and MSHA approvals)
  703+ATEX, 705+ATEX, and 706RC+ATEX (ATEX approvals).
Any questions regarding our I.S. products (Spark / Spartan), ask before making changes or when other issues arise
  Intrinsic Safety Product (ISP) Manager – Alex Blodgett
  Backup ISP Manager – James Higley
D0001.0014 Quality Addendum for I.S. Product details quality requirements for I.S. products, please review the latest version
D0001.0014-1 Intrinsic Safety Listings and Approvals details approvals and certificates for our I.S. products, please review the latest version

## Slide 3: Larson Davis Product Area Classifications / Markings
  Spartan 730IS Dosimeter:
    Class I, Zone 0, AEx ia IIC T165°C Ga
    II 1 G Ex ia IIC T165°C Ga
    I 1 M Ex ia I Ma
    Class I, Division 1, Groups A, B, C & D, T165°C
  Spark Noise Dosimeters:
    Zone 1, II 2G Ex ib IIB T4…T3 Gb
    Class I,II,III Division 1 Groups C,D,E,F,G T4
    MSHA (Has its own approval)

## Slide 4: Schedule Drawings / I.S. Part Numbering
  Schedule Drawings
    Controlled drawings in Intrinsic Safety Listing / Approvals
    Cannot modify without approval body review ($$$)
    Special wording on each page of drawing: “Do not change drawing without approval from the notified body”
    Drawing number will have “-IS-SD” suffix
    Must be followed completely
  Related Drawings
    Provide full details for manufacturing IS Product
    Changes not affecting Schedule drawing may be made without approval body (after ENG /ISP Manager review)
    Drawing number have “-IS” suffix on drawing / part number
    Contain special wording linking to Schedule drawing and warning changes must be approved by ISP Manager/ENG
    Other manufacturer’s document numbers also have “-IS” suffix

## Slide 5: More Resources
D0001.0014 Quality Addendum for Intrinsic Safety Product
D0001.0014-1 Intrinsic Safety Listing and Approvals
More in depth Hazardous Location training available on the network
  R:\Provo\Shared\ATEX - Explosive Atmospheres\ATEX Training\
    Larson Davis Intrinsic Safety Training New Employee 2020
    PC Basic Hazloc Testing and Evaluation - Self Study Training
    Intertek Guide to Hazardous Location Equipment Marking
Hazardous Location Explosion Videos on network:
  R:\Provo\Shared\ATEX - Explosive Atmospheres\ATEX Training\Hazardous Location Explosion Videos
Ask ISP Manager if you have any questions