SHIPPING HVM100 FAMILY

1.0		APPLICABILITY

This instruction applies to the HVM100, IHVM100 and HAVPro products and provides the instruction for preparing the products for shipping.

2.0		RESPONSIBILITY & AUTHORITY

The Technician or individual assigned to prepare the product for shipping has the responsibility to follow this instruction.

3.0		HVM100 / HAVPRO SHIPPING PROCESS

- Check calibration dates on Certification documents and Cal tag, and Cal seal attached.

- Retest the unit if found to be out of the two month shelf calibration period.

- On the HVM final checklist, list the Order # and Calibration date.

- Determine the Languages and Firmware Options needed and mark on the final checklist.

- Connect the unit to the computer using the CBL006 cable, and power adapter.

- Turn the unit on, and push the “Tools” key, then up arrow to the Baud rate. Press the center “” check key, and arrow to 115.2k rate, and push the “” check key again to set.

- Open Larson Davis HVM Firmware Upgrade Wizard. Select “next” to begin.

- Program will start STEP 1. “” Automatically detect communication settings and select NEXT to continue.

- Select language pack option: L1, L2, L3 or L4 and select FINISH.

- HVMProgrammer warning opens and asks if you want to continue, select Yes. Bootcode is programmed.

- Power-cycle opens. Turn the instrument off and back on. “” Instrument has been power-cycled. Select OK.

- Programming firmware begins. Firmware successful, select OK

- Verify firmware version is correct.

- Turn unit off and back on

- Verify that the correct languages have been loaded, push tools key and arrow up three times, press “” check key and arrow through languages stopping on English again and pressing “” check key to enter. The list below shows which languages are loaded in each version.

L1:		English, Espanol, Deutsch, Italiano

L2:		English, Deutsch, Portuguese, Cesky

L3:		English, Francais, Italiano, Deutsch

L4:		English, Nederland’s, Deutsch, Francais

3.1		Enable Appropriate Options

- Open HVM Configuration Tool Excel File

- Enter serial number and selected options under Serial Number and Options Command Generator

  - < Enter serial number

  - < Enable 100 file option? (0 or 1) This is always marked 1

  - < Enable Hand-Arm option? (0 or 1) Change to 1 to enable option, leave at 0 if option not wanted

  - < Enable Whole Body option? (0 or 1) Change to 1 to enable option, leave at 0 if option not wanted

- Red text under Command will change depending on options selected above.

- Record option codes (as per table) on the HVM Final checklist.

| Option | Send |  |  |  |  | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| No option specified |  | #2 |  |  | #5 | all orders get at least the #2 and # 5 |
| -HA, -GT, or -HAGT |  | #2 | #3 |  | #5 |  |
| -WB |  | #2 |  | #4 | #5 |  |
| -ALL, or -HAWB |  | #2 | #3 | #4 | #5 |  |

- Right-click and copy the option codes (red text) in yellow highlighted cells.

Make sure instrument is turned on

- Open Tera Term – “Ttermpro shortcut”

- Right click and paste copied option codes

- Select OK and Close

3.2		Verify Options

1. Verify that the appropriate options have been enabled.

  - Hit the “Set Up” key and then the check key, verify the options by using the arrow key.

  - Go back to “Vibration” and push the “” check key.

  - Hit “Data” key then the “Recall” key. Arrow to the left and check for 50 + files.

  - Turn Off unit

      - Store set-up? Select NO

3.3		Clear All Files and Setups

- Erase all Files and Setups,

- Press “Tools” key and arrow up six times, push “” check key and arrow to YES and press “” check key.

Arrow up one more time to Erase all files and repeat step to erase.

- Enter Order information into Larson-Davis database.  The HVM100 and HAVPro are each considered unique instruments in the LD database.

3.4		Final Inspection

Performed by QC

- Install 2 AA batteries and turn on the DUT.

- Verify that the serial number is correct.

- Turn off the DUT and remove the batteries.

- Inspect all parts for visible defects and verify that the assembled DUT includes the following parts.

  - Case screws (5)

  - Large rubber foot

  - Small rubber feet (2)

  - Serial number label – verify that serial number on the label is correct

  - Calibration label – NOTE, the initial calibration period for the HAVPro is 15 months.

  - Battery door – batteries should be removed

  - Jewel

  - Camera mount insert (helicoil)

4.0		Records & Distribution

No additional records are generated as a result of this instruction.  Availability of the instruction is through the online area of Document Control.

5.0		Revision History

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 975 | B | 9/12/06 | RH | Revised & simplified |
| 1751 | C | 10/16/17 | AW | Updated instruction to better fit current procedures |