- Field Instrument Service

1.0	PURPOSE AND SCOPE

To provide a procedure to service, troubleshoot, and test instruments away from the factory.  The procedure applies to customer instruments, and marketing demonstration equipment in need of repair or calibration.

2.0	AFFECTED DEPARTMENTS

- Repair & Calibration

- Marketing & Sales

- Quality

3.0	REFERENCE DOCUMENTS

- Troubleshooting checklist for instrument needing repair.

- Metrological System Controls & Processes

4.0	RESPONSIBILITIES & AUTHORITY

The Repair and Calibration Manager has the responsibility to verify that pre-approval for the onsite service has been granted prior to making travel arrangements.

The Repair and Calibration Representative has the following responsibilities and authority:

- Locate the equipment.

- Diagnose equipment status.

- Determine a manufacture’s defect or warranty period.

- Contact customer with estimated quote for costs and completion date.

- Obtain PO or billing instructions from the customer before work is performed.

- Contact Larson-Davis for additional parts and/or supplies

- Repair, if needed

- Test functionality and reliability.

5.0	DEFINITIONS

- LD:	Larson-Davis

- PO:	Purchase Order

- WO:	Work Order

6.0	SAFETY PRECAUTIONS

- Wear appropriate eyewear for troubleshooting and repairing instruments.

- Observe safe lifting practices.  If object is too heavy or awkward, get help to lift it.

- If working in areas with the potential for head injury (low hanging branches, construction, etc.) wear a hard hat.

- If working from scaffolding or a ladder, make sure the platform is secure and will not move.  If appropriate use an OSHA approved body harness for fall protection.

- Take all necessary electrical precautions for lock-out / tag-out and for working in areas with overhead power lines.

- If injured while onsite, obtain medical help then contact Larson-Davis and report the incident.

7.0	EQUIPMENT & MATERIALS

- PC with test programs

- Tools for on-site repair

- Spare parts as applicable

8.0	INSTRUCTIONS

8.1	Prior to Going Onsite

Collect the following information or materials as needed for the type of instrument and/or installation.

- Locate any prior field installation notes for the instrument or the site.

- Date of last calibration or service.

- Is this a repair or a routine calibration with maintenance?

- When did the product fail?

- What type of environment is the instrument being used in?

- What were the symptoms of the failure?

- Has anything recently changed in the area?

- Determine if any additional equipment or parts are needed to complete the job.

8.2	Onsite

Once onsite, the Representative will review the information, collected in Section 8.1, and evaluate the status of the instrument or installation.

- Has there been any change in the area since the last visit?

- Go through the "trouble checklist" for the instrument.

- Reset Memory.

- Exchange modules to identify the fault.

- List any suggested improvements.

- If a problem is found, how can the problem be avoided in the future?

9.0	INSPECTION

- Check the equipment for signs of damage.

- Check the equipment operation over time.

- If an outdoor system, verify that locks work and that box is secure.

- Verify batteries, if applicable, are still functioning.

10.0	RECORDS

Unless otherwise stated, the records are maintained for at least one year.  They may be maintained in hardcopy or electronic format by the Calibration & Repair Department.

- Work Order for service

- Shipping document

- Warranty status

- Error codes

- Expense Records

11.0	DISTRIBUTION

- Calibration & Repair Department

- Marketing & Sales

12.0	ATTACHMENTS

- Trouble checklist.

- Completed service work order.

- Customer acknowledgement form.

13.0	REVISION HISTORY

| ECO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |