### HVM100 FINAL ASSEMBLY AND TEST

1.0	PURPOSE AND SCOPE

This document describes the test procedure for verifying that the HVM100, IHVM100, or HAVPro is functioning within specified parameters.  All employees who have responsibility for testing this component are required the follow the instructions detailed in this procedure.

2.0	AFFECTED DEPARTMENTS

- Manufacturing

3.0	REFERENCE DOCUMENTS

- Current revision assembly drawings for the HVM100, IHVM100, or HAVPro

- Device under test User’s Manual (IHVM100.01, IIHVM100.01, or IHAVPRO.01)

- HVM100 FINAL CHECKLIST (see D0001.8099 HVM100 FINAL INSPECTION)

4.0	RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specifications.

- Troubleshoot and correct product as required.

- Communicate concerns to the Supervisor of Quality Assurance.

- Request management review of product concerns.

- Follow established ESD standards.

5.0	DEFINITIONS

- The tests in this document are designed to verify that an HVM100, IHVM100, or HAVPro instrument is functioning properly.  This document refers to the instrument being tested as the device under test (DUT).

6.0	SAFETY PRECAUTIONS

- Safety glasses when soldering, lead clipping, or testing power supplies.

7.0	EQUIPMENT AND MATERIALS

## Phillips Screwdriver

- External Power Adapter (i.e. PSA027 or equivalent)

## Serial Communications Cable (CBL006)

- CBL106 Cable	(HVM100 Direct Input Test Cable, Lemo to BNC)

- CBL107 Cable	(HVM100 Charge Input Test Cable, Lemo to BNC, 1000pF)

- LD# CBL121 Cable (HVM100 INPUT - 3 X BNC  MALE TO LEMO 5')

- Computer (PC 486 66MHz or newer, running Windows 95 or newer)

- HVMTest.exe Windows software installed on the test computer

- Larson Davis Test Station (2900, 2209, Computer)

- Window’s Hyperterminal or equivalent

- 2900 Signal Generator Control Software installed on the test computer (SG.exe)

- Bench-top Power Supply

8.0	INSTRUCTIONS

##### NOTE:  Some of the following instructions involve setting parameters on the DUT or reading data from the DUT.  Consult the DUT user’s manual (IHVM100.01, IIHVM100.01, or IHAVPRO.01) for detailed operating instructions  .

8.1	Assemble the DUT and Load the Serial Number

- Assemble the DUT according to the assembly drawing specified by the bill of manufacturing.

- Connect an external power adapter (PSA027) to the DUT and turn on the DUT.

- Connect the CBL006 cable to the DUT serial connector and to the PC’s serial port.

- Setup Window’s Hyperterminal for 115.2k baud, 8 data bits, 1 stop bit, no parity.  Using Windows Hyperterminal, send an R1 command to the DUT.  The DUT will respond with one of the following messages (based on which type of instrument is being assembled).

##### HVM100:		LARSON DAVIS HVM100 65535 rev x.xx

IHVM100:		LARSON DAVIS IHVM100 65535 rev x.xx

HAVPro:		QUEST TECHNOLOGIES HAVPro 65545 rev x.xx

Where x.xx is the current firmware version.

- Load the next available serial number in the DUT by sending a #46327,aaa command followed by a #57391,bbb command.  The formula below shows how to calculate the aaa and bbb values based on the serial number.

aaa is the remainder of the Serial Number divided by 256

bbb is the whole number result of the Serial Number divided by 256

For example, if the Serial Number is 00167, then 167/256 equals 0 with a remainder of 167; therefore, aaa = 167 and bbb = 0.

If the Serial Number is 00731, then 731/256 eauals 2 with a remainder of 219;  therefore, aaa = 167 and bbb = 2.

The tables on the following pages show the appropriate aaa and bbb values for each serial number.

- NOTE:  The HVM100, IHVM100, and HAVPro all use the same type of 5 digit serial number.  Therefore it is possible for an HVM100, IHVM100, and HAVPro to all have the same serial number.  However, no two HVM100s, or two IHVM100s, or two HAVPros should have identical serial numbers.  Be sure to assign the next appropriate serial number based on which type of instrument is being assembled.

- The HAVPro serial numbers also include a date of manufacture code.  For example, units manufactured in 2003 are to be serialized as 03001, 03002, 03003, etc.  Units manufactured in 2004 will be 04001, 04002, 04003, and so forth for subsequent years.   Since the HAVPro serialization scheme is somewhat different, separate tables for the HAVPro are shown below.

- Turn the DUT off and then back on again to verify that the correct serial number has been loaded (the DUT displays its serial number at power-up).

Send

#46327,aaa			HVM100 and IHVM100:  SERIAL NUMBERS 100 thru 299

followed by

#57391,bbb

| SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 100 | 100 | 0 | 150 | 150 | 0 | 200 | 200 | 0 | 250 | 250 | 0 |
| 101 | 101 | 0 | 151 | 151 | 0 | 201 | 201 | 0 | 251 | 251 | 0 |
| 102 | 102 | 0 | 152 | 152 | 0 | 202 | 202 | 0 | 252 | 252 | 0 |
| 103 | 103 | 0 | 153 | 153 | 0 | 203 | 203 | 0 | 253 | 253 | 0 |
| 104 | 104 | 0 | 154 | 154 | 0 | 204 | 204 | 0 | 254 | 254 | 0 |
| 105 | 105 | 0 | 155 | 155 | 0 | 205 | 205 | 0 | 255 | 255 | 0 |
| 106 | 106 | 0 | 156 | 156 | 0 | 206 | 206 | 0 | 256 | 0 | 1 |
| 107 | 107 | 0 | 157 | 157 | 0 | 207 | 207 | 0 | 257 | 1 | 1 |
| 108 | 108 | 0 | 158 | 158 | 0 | 208 | 208 | 0 | 258 | 2 | 1 |
| 109 | 109 | 0 | 159 | 159 | 0 | 209 | 209 | 0 | 259 | 3 | 1 |
| 110 | 110 | 0 | 160 | 160 | 0 | 210 | 210 | 0 | 260 | 4 | 1 |
| 111 | 111 | 0 | 161 | 161 | 0 | 211 | 211 | 0 | 261 | 5 | 1 |
| 112 | 112 | 0 | 162 | 162 | 0 | 212 | 212 | 0 | 262 | 6 | 1 |
| 113 | 113 | 0 | 163 | 163 | 0 | 213 | 213 | 0 | 263 | 7 | 1 |
| 114 | 114 | 0 | 164 | 164 | 0 | 214 | 214 | 0 | 264 | 8 | 1 |
| 115 | 115 | 0 | 165 | 165 | 0 | 215 | 215 | 0 | 265 | 9 | 1 |
| 116 | 116 | 0 | 166 | 166 | 0 | 216 | 216 | 0 | 266 | 10 | 1 |
| 117 | 117 | 0 | 167 | 167 | 0 | 217 | 217 | 0 | 267 | 11 | 1 |
| 118 | 118 | 0 | 168 | 168 | 0 | 218 | 218 | 0 | 268 | 12 | 1 |
| 119 | 119 | 0 | 169 | 169 | 0 | 219 | 219 | 0 | 269 | 13 | 1 |
| 120 | 120 | 0 | 170 | 170 | 0 | 220 | 220 | 0 | 270 | 14 | 1 |
| 121 | 121 | 0 | 171 | 171 | 0 | 221 | 221 | 0 | 271 | 15 | 1 |
| 122 | 122 | 0 | 172 | 172 | 0 | 222 | 222 | 0 | 272 | 16 | 1 |
| 123 | 123 | 0 | 173 | 173 | 0 | 223 | 223 | 0 | 273 | 17 | 1 |
| 124 | 124 | 0 | 174 | 174 | 0 | 224 | 224 | 0 | 274 | 18 | 1 |
| 125 | 125 | 0 | 175 | 175 | 0 | 225 | 225 | 0 | 275 | 19 | 1 |
| 126 | 126 | 0 | 176 | 176 | 0 | 226 | 226 | 0 | 276 | 20 | 1 |
| 127 | 127 | 0 | 177 | 177 | 0 | 227 | 227 | 0 | 277 | 21 | 1 |
| 128 | 128 | 0 | 178 | 178 | 0 | 228 | 228 | 0 | 278 | 22 | 1 |
| 129 | 129 | 0 | 179 | 179 | 0 | 229 | 229 | 0 | 279 | 23 | 1 |
| 130 | 130 | 0 | 180 | 180 | 0 | 230 | 230 | 0 | 280 | 24 | 1 |
| 131 | 131 | 0 | 181 | 181 | 0 | 231 | 231 | 0 | 281 | 25 | 1 |
| 132 | 132 | 0 | 182 | 182 | 0 | 232 | 232 | 0 | 282 | 26 | 1 |
| 133 | 133 | 0 | 183 | 183 | 0 | 233 | 233 | 0 | 283 | 27 | 1 |
| 134 | 134 | 0 | 184 | 184 | 0 | 234 | 234 | 0 | 284 | 28 | 1 |
| 135 | 135 | 0 | 185 | 185 | 0 | 235 | 235 | 0 | 285 | 29 | 1 |
| 136 | 136 | 0 | 186 | 186 | 0 | 236 | 236 | 0 | 286 | 30 | 1 |
| 137 | 137 | 0 | 187 | 187 | 0 | 237 | 237 | 0 | 287 | 31 | 1 |
| 138 | 138 | 0 | 188 | 188 | 0 | 238 | 238 | 0 | 288 | 32 | 1 |
| 139 | 139 | 0 | 189 | 189 | 0 | 239 | 239 | 0 | 289 | 33 | 1 |
| 140 | 140 | 0 | 190 | 190 | 0 | 240 | 240 | 0 | 290 | 34 | 1 |
| 141 | 141 | 0 | 191 | 191 | 0 | 241 | 241 | 0 | 291 | 35 | 1 |
| 142 | 142 | 0 | 192 | 192 | 0 | 242 | 242 | 0 | 292 | 36 | 1 |
| 143 | 143 | 0 | 193 | 193 | 0 | 243 | 243 | 0 | 293 | 37 | 1 |
| 144 | 144 | 0 | 194 | 194 | 0 | 244 | 244 | 0 | 294 | 38 | 1 |
| 145 | 145 | 0 | 195 | 195 | 0 | 245 | 245 | 0 | 295 | 39 | 1 |
| 146 | 146 | 0 | 196 | 196 | 0 | 246 | 246 | 0 | 296 | 40 | 1 |
| 147 | 147 | 0 | 197 | 197 | 0 | 247 | 247 | 0 | 297 | 41 | 1 |
| 148 | 148 | 0 | 198 | 198 | 0 | 248 | 248 | 0 | 298 | 42 | 1 |
| 149 | 149 | 0 | 199 | 199 | 0 | 249 | 249 | 0 | 299 | 43 | 1 |

Send

#46327,aaa			HVM100 and IHVM100:  SERIAL NUMBERS 300 thru 499

followed by

#57391,bbb

| SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 300 | 44 | 1 | 350 | 94 | 1 | 400 | 144 | 1 | 450 | 194 | 1 |
| 301 | 45 | 1 | 351 | 95 | 1 | 401 | 145 | 1 | 451 | 195 | 1 |
| 302 | 46 | 1 | 352 | 96 | 1 | 402 | 146 | 1 | 452 | 196 | 1 |
| 303 | 47 | 1 | 353 | 97 | 1 | 403 | 147 | 1 | 453 | 197 | 1 |
| 304 | 48 | 1 | 354 | 98 | 1 | 404 | 148 | 1 | 454 | 198 | 1 |
| 305 | 49 | 1 | 355 | 99 | 1 | 405 | 149 | 1 | 455 | 199 | 1 |
| 306 | 50 | 1 | 356 | 100 | 1 | 406 | 150 | 1 | 456 | 200 | 1 |
| 307 | 51 | 1 | 357 | 101 | 1 | 407 | 151 | 1 | 457 | 201 | 1 |
| 308 | 52 | 1 | 358 | 102 | 1 | 408 | 152 | 1 | 458 | 202 | 1 |
| 309 | 53 | 1 | 359 | 103 | 1 | 409 | 153 | 1 | 459 | 203 | 1 |
| 310 | 54 | 1 | 360 | 104 | 1 | 410 | 154 | 1 | 460 | 204 | 1 |
| 311 | 55 | 1 | 361 | 105 | 1 | 411 | 155 | 1 | 461 | 205 | 1 |
| 312 | 56 | 1 | 362 | 106 | 1 | 412 | 156 | 1 | 462 | 206 | 1 |
| 313 | 57 | 1 | 363 | 107 | 1 | 413 | 157 | 1 | 463 | 207 | 1 |
| 314 | 58 | 1 | 364 | 108 | 1 | 414 | 158 | 1 | 464 | 208 | 1 |
| 315 | 59 | 1 | 365 | 109 | 1 | 415 | 159 | 1 | 465 | 209 | 1 |
| 316 | 60 | 1 | 366 | 110 | 1 | 416 | 160 | 1 | 466 | 210 | 1 |
| 317 | 61 | 1 | 367 | 111 | 1 | 417 | 161 | 1 | 467 | 211 | 1 |
| 318 | 62 | 1 | 368 | 112 | 1 | 418 | 162 | 1 | 468 | 212 | 1 |
| 319 | 63 | 1 | 369 | 113 | 1 | 419 | 163 | 1 | 469 | 213 | 1 |
| 320 | 64 | 1 | 370 | 114 | 1 | 420 | 164 | 1 | 470 | 214 | 1 |
| 321 | 65 | 1 | 371 | 115 | 1 | 421 | 165 | 1 | 471 | 215 | 1 |
| 322 | 66 | 1 | 372 | 116 | 1 | 422 | 166 | 1 | 472 | 216 | 1 |
| 323 | 67 | 1 | 373 | 117 | 1 | 423 | 167 | 1 | 473 | 217 | 1 |
| 324 | 68 | 1 | 374 | 118 | 1 | 424 | 168 | 1 | 474 | 218 | 1 |
| 325 | 69 | 1 | 375 | 119 | 1 | 425 | 169 | 1 | 475 | 219 | 1 |
| 326 | 70 | 1 | 376 | 120 | 1 | 426 | 170 | 1 | 476 | 220 | 1 |
| 327 | 71 | 1 | 377 | 121 | 1 | 427 | 171 | 1 | 477 | 221 | 1 |
| 328 | 72 | 1 | 378 | 122 | 1 | 428 | 172 | 1 | 478 | 222 | 1 |
| 329 | 73 | 1 | 379 | 123 | 1 | 429 | 173 | 1 | 479 | 223 | 1 |
| 330 | 74 | 1 | 380 | 124 | 1 | 430 | 174 | 1 | 480 | 224 | 1 |
| 331 | 75 | 1 | 381 | 125 | 1 | 431 | 175 | 1 | 481 | 225 | 1 |
| 332 | 76 | 1 | 382 | 126 | 1 | 432 | 176 | 1 | 482 | 226 | 1 |
| 333 | 77 | 1 | 383 | 127 | 1 | 433 | 177 | 1 | 483 | 227 | 1 |
| 334 | 78 | 1 | 384 | 128 | 1 | 434 | 178 | 1 | 484 | 228 | 1 |
| 335 | 79 | 1 | 385 | 129 | 1 | 435 | 179 | 1 | 485 | 229 | 1 |
| 336 | 80 | 1 | 386 | 130 | 1 | 436 | 180 | 1 | 486 | 230 | 1 |
| 337 | 81 | 1 | 387 | 131 | 1 | 437 | 181 | 1 | 487 | 231 | 1 |
| 338 | 82 | 1 | 388 | 132 | 1 | 438 | 182 | 1 | 488 | 232 | 1 |
| 339 | 83 | 1 | 389 | 133 | 1 | 439 | 183 | 1 | 489 | 233 | 1 |
| 340 | 84 | 1 | 390 | 134 | 1 | 440 | 184 | 1 | 490 | 234 | 1 |
| 341 | 85 | 1 | 391 | 135 | 1 | 441 | 185 | 1 | 491 | 235 | 1 |
| 342 | 86 | 1 | 392 | 136 | 1 | 442 | 186 | 1 | 492 | 236 | 1 |
| 343 | 87 | 1 | 393 | 137 | 1 | 443 | 187 | 1 | 493 | 237 | 1 |
| 344 | 88 | 1 | 394 | 138 | 1 | 444 | 188 | 1 | 494 | 238 | 1 |
| 345 | 89 | 1 | 395 | 139 | 1 | 445 | 189 | 1 | 495 | 239 | 1 |
| 346 | 90 | 1 | 396 | 140 | 1 | 446 | 190 | 1 | 496 | 240 | 1 |
| 347 | 91 | 1 | 397 | 141 | 1 | 447 | 191 | 1 | 497 | 241 | 1 |
| 348 | 92 | 1 | 398 | 142 | 1 | 448 | 192 | 1 | 498 | 242 | 1 |
| 349 | 93 | 1 | 399 | 143 | 1 | 449 | 193 | 1 | 499 | 243 | 1 |

Send

#46327,aaa			HVM100 and IHVM100:  SERIAL NUMBERS 500 thru 699

followed by

#57391,bbb

| SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 500 | 244 | 1 | 550 | 38 | 2 | 600 | 88 | 2 | 650 | 138 | 2 |
| 501 | 245 | 1 | 551 | 39 | 2 | 601 | 89 | 2 | 651 | 139 | 2 |
| 502 | 246 | 1 | 552 | 40 | 2 | 602 | 90 | 2 | 652 | 140 | 2 |
| 503 | 247 | 1 | 553 | 41 | 2 | 603 | 91 | 2 | 653 | 141 | 2 |
| 504 | 248 | 1 | 554 | 42 | 2 | 604 | 92 | 2 | 654 | 142 | 2 |
| 505 | 249 | 1 | 555 | 43 | 2 | 605 | 93 | 2 | 655 | 143 | 2 |
| 506 | 250 | 1 | 556 | 44 | 2 | 606 | 94 | 2 | 656 | 144 | 2 |
| 507 | 251 | 1 | 557 | 45 | 2 | 607 | 95 | 2 | 657 | 145 | 2 |
| 508 | 252 | 1 | 558 | 46 | 2 | 608 | 96 | 2 | 658 | 146 | 2 |
| 509 | 253 | 1 | 559 | 47 | 2 | 609 | 97 | 2 | 659 | 147 | 2 |
| 510 | 254 | 1 | 560 | 48 | 2 | 610 | 98 | 2 | 660 | 148 | 2 |
| 511 | 255 | 1 | 561 | 49 | 2 | 611 | 99 | 2 | 661 | 149 | 2 |
| 512 | 0 | 2 | 562 | 50 | 2 | 612 | 100 | 2 | 662 | 150 | 2 |
| 513 | 1 | 2 | 563 | 51 | 2 | 613 | 101 | 2 | 663 | 151 | 2 |
| 514 | 2 | 2 | 564 | 52 | 2 | 614 | 102 | 2 | 664 | 152 | 2 |
| 515 | 3 | 2 | 565 | 53 | 2 | 615 | 103 | 2 | 665 | 153 | 2 |
| 516 | 4 | 2 | 566 | 54 | 2 | 616 | 104 | 2 | 666 | 154 | 2 |
| 517 | 5 | 2 | 567 | 55 | 2 | 617 | 105 | 2 | 667 | 155 | 2 |
| 518 | 6 | 2 | 568 | 56 | 2 | 618 | 106 | 2 | 668 | 156 | 2 |
| 519 | 7 | 2 | 569 | 57 | 2 | 619 | 107 | 2 | 669 | 157 | 2 |
| 520 | 8 | 2 | 570 | 58 | 2 | 620 | 108 | 2 | 670 | 158 | 2 |
| 521 | 9 | 2 | 571 | 59 | 2 | 621 | 109 | 2 | 671 | 159 | 2 |
| 522 | 10 | 2 | 572 | 60 | 2 | 622 | 110 | 2 | 672 | 160 | 2 |
| 523 | 11 | 2 | 573 | 61 | 2 | 623 | 111 | 2 | 673 | 161 | 2 |
| 524 | 12 | 2 | 574 | 62 | 2 | 624 | 112 | 2 | 674 | 162 | 2 |
| 525 | 13 | 2 | 575 | 63 | 2 | 625 | 113 | 2 | 675 | 163 | 2 |
| 526 | 14 | 2 | 576 | 64 | 2 | 626 | 114 | 2 | 676 | 164 | 2 |
| 527 | 15 | 2 | 577 | 65 | 2 | 627 | 115 | 2 | 677 | 165 | 2 |
| 528 | 16 | 2 | 578 | 66 | 2 | 628 | 116 | 2 | 678 | 166 | 2 |
| 529 | 17 | 2 | 579 | 67 | 2 | 629 | 117 | 2 | 679 | 167 | 2 |
| 530 | 18 | 2 | 580 | 68 | 2 | 630 | 118 | 2 | 680 | 168 | 2 |
| 531 | 19 | 2 | 581 | 69 | 2 | 631 | 119 | 2 | 681 | 169 | 2 |
| 532 | 20 | 2 | 582 | 70 | 2 | 632 | 120 | 2 | 682 | 170 | 2 |
| 533 | 21 | 2 | 583 | 71 | 2 | 633 | 121 | 2 | 683 | 171 | 2 |
| 534 | 22 | 2 | 584 | 72 | 2 | 634 | 122 | 2 | 684 | 172 | 2 |
| 535 | 23 | 2 | 585 | 73 | 2 | 635 | 123 | 2 | 685 | 173 | 2 |
| 536 | 24 | 2 | 586 | 74 | 2 | 636 | 124 | 2 | 686 | 174 | 2 |
| 537 | 25 | 2 | 587 | 75 | 2 | 637 | 125 | 2 | 687 | 175 | 2 |
| 538 | 26 | 2 | 588 | 76 | 2 | 638 | 126 | 2 | 688 | 176 | 2 |
| 539 | 27 | 2 | 589 | 77 | 2 | 639 | 127 | 2 | 689 | 177 | 2 |
| 540 | 28 | 2 | 590 | 78 | 2 | 640 | 128 | 2 | 690 | 178 | 2 |
| 541 | 29 | 2 | 591 | 79 | 2 | 641 | 129 | 2 | 691 | 179 | 2 |
| 542 | 30 | 2 | 592 | 80 | 2 | 642 | 130 | 2 | 692 | 180 | 2 |
| 543 | 31 | 2 | 593 | 81 | 2 | 643 | 131 | 2 | 693 | 181 | 2 |
| 544 | 32 | 2 | 594 | 82 | 2 | 644 | 132 | 2 | 694 | 182 | 2 |
| 545 | 33 | 2 | 595 | 83 | 2 | 645 | 133 | 2 | 695 | 183 | 2 |
| 546 | 34 | 2 | 596 | 84 | 2 | 646 | 134 | 2 | 696 | 184 | 2 |
| 547 | 35 | 2 | 597 | 85 | 2 | 647 | 135 | 2 | 697 | 185 | 2 |
| 548 | 36 | 2 | 598 | 86 | 2 | 648 | 136 | 2 | 698 | 186 | 2 |
| 549 | 37 | 2 | 599 | 87 | 2 | 649 | 137 | 2 | 699 | 187 | 2 |

Send

#46327,aaa			HVM100 and IHVM100:  SERIAL NUMBERS 700 thru 899

followed by

#57391,bbb

| SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 700 | 188 | 2 | 750 | 238 | 2 | 800 | 32 | 3 | 850 | 82 | 3 |
| 701 | 189 | 2 | 751 | 239 | 2 | 801 | 33 | 3 | 851 | 83 | 3 |
| 702 | 190 | 2 | 752 | 240 | 2 | 802 | 34 | 3 | 852 | 84 | 3 |
| 703 | 191 | 2 | 753 | 241 | 2 | 803 | 35 | 3 | 853 | 85 | 3 |
| 704 | 192 | 2 | 754 | 242 | 2 | 804 | 36 | 3 | 854 | 86 | 3 |
| 705 | 193 | 2 | 755 | 243 | 2 | 805 | 37 | 3 | 855 | 87 | 3 |
| 706 | 194 | 2 | 756 | 244 | 2 | 806 | 38 | 3 | 856 | 88 | 3 |
| 707 | 195 | 2 | 757 | 245 | 2 | 807 | 39 | 3 | 857 | 89 | 3 |
| 708 | 196 | 2 | 758 | 246 | 2 | 808 | 40 | 3 | 858 | 90 | 3 |
| 709 | 197 | 2 | 759 | 247 | 2 | 809 | 41 | 3 | 859 | 91 | 3 |
| 710 | 198 | 2 | 760 | 248 | 2 | 810 | 42 | 3 | 860 | 92 | 3 |
| 711 | 199 | 2 | 761 | 249 | 2 | 811 | 43 | 3 | 861 | 93 | 3 |
| 712 | 200 | 2 | 762 | 250 | 2 | 812 | 44 | 3 | 862 | 94 | 3 |
| 713 | 201 | 2 | 763 | 251 | 2 | 813 | 45 | 3 | 863 | 95 | 3 |
| 714 | 202 | 2 | 764 | 252 | 2 | 814 | 46 | 3 | 864 | 96 | 3 |
| 715 | 203 | 2 | 765 | 253 | 2 | 815 | 47 | 3 | 865 | 97 | 3 |
| 716 | 204 | 2 | 766 | 254 | 2 | 816 | 48 | 3 | 866 | 98 | 3 |
| 717 | 205 | 2 | 767 | 255 | 2 | 817 | 49 | 3 | 867 | 99 | 3 |
| 718 | 206 | 2 | 768 | 0 | 3 | 818 | 50 | 3 | 868 | 100 | 3 |
| 719 | 207 | 2 | 769 | 1 | 3 | 819 | 51 | 3 | 869 | 101 | 3 |
| 720 | 208 | 2 | 770 | 2 | 3 | 820 | 52 | 3 | 870 | 102 | 3 |
| 721 | 209 | 2 | 771 | 3 | 3 | 821 | 53 | 3 | 871 | 103 | 3 |
| 722 | 210 | 2 | 772 | 4 | 3 | 822 | 54 | 3 | 872 | 104 | 3 |
| 723 | 211 | 2 | 773 | 5 | 3 | 823 | 55 | 3 | 873 | 105 | 3 |
| 724 | 212 | 2 | 774 | 6 | 3 | 824 | 56 | 3 | 874 | 106 | 3 |
| 725 | 213 | 2 | 775 | 7 | 3 | 825 | 57 | 3 | 875 | 107 | 3 |
| 726 | 214 | 2 | 776 | 8 | 3 | 826 | 58 | 3 | 876 | 108 | 3 |
| 727 | 215 | 2 | 777 | 9 | 3 | 827 | 59 | 3 | 877 | 109 | 3 |
| 728 | 216 | 2 | 778 | 10 | 3 | 828 | 60 | 3 | 878 | 110 | 3 |
| 729 | 217 | 2 | 779 | 11 | 3 | 829 | 61 | 3 | 879 | 111 | 3 |
| 730 | 218 | 2 | 780 | 12 | 3 | 830 | 62 | 3 | 880 | 112 | 3 |
| 731 | 219 | 2 | 781 | 13 | 3 | 831 | 63 | 3 | 881 | 113 | 3 |
| 732 | 220 | 2 | 782 | 14 | 3 | 832 | 64 | 3 | 882 | 114 | 3 |
| 733 | 221 | 2 | 783 | 15 | 3 | 833 | 65 | 3 | 883 | 115 | 3 |
| 734 | 222 | 2 | 784 | 16 | 3 | 834 | 66 | 3 | 884 | 116 | 3 |
| 735 | 223 | 2 | 785 | 17 | 3 | 835 | 67 | 3 | 885 | 117 | 3 |
| 736 | 224 | 2 | 786 | 18 | 3 | 836 | 68 | 3 | 886 | 118 | 3 |
| 737 | 225 | 2 | 787 | 19 | 3 | 837 | 69 | 3 | 887 | 119 | 3 |
| 738 | 226 | 2 | 788 | 20 | 3 | 838 | 70 | 3 | 888 | 120 | 3 |
| 739 | 227 | 2 | 789 | 21 | 3 | 839 | 71 | 3 | 889 | 121 | 3 |
| 740 | 228 | 2 | 790 | 22 | 3 | 840 | 72 | 3 | 890 | 122 | 3 |
| 741 | 229 | 2 | 791 | 23 | 3 | 841 | 73 | 3 | 891 | 123 | 3 |
| 742 | 230 | 2 | 792 | 24 | 3 | 842 | 74 | 3 | 892 | 124 | 3 |
| 743 | 231 | 2 | 793 | 25 | 3 | 843 | 75 | 3 | 893 | 125 | 3 |
| 744 | 232 | 2 | 794 | 26 | 3 | 844 | 76 | 3 | 894 | 126 | 3 |
| 745 | 233 | 2 | 795 | 27 | 3 | 845 | 77 | 3 | 895 | 127 | 3 |
| 746 | 234 | 2 | 796 | 28 | 3 | 846 | 78 | 3 | 896 | 128 | 3 |
| 747 | 235 | 2 | 797 | 29 | 3 | 847 | 79 | 3 | 897 | 129 | 3 |
| 748 | 236 | 2 | 798 | 30 | 3 | 848 | 80 | 3 | 898 | 130 | 3 |
| 749 | 237 | 2 | 799 | 31 | 3 | 849 | 81 | 3 | 899 | 131 | 3 |

Send

#46327,aaa			HVM100 and IHVM100:  SERIAL NUMBERS 900 thru 999

followed by

#57391,bbb

| SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 900 | 132 | 3 | 950 | 182 | 3 |  |  |  |  |  |  |
| 901 | 133 | 3 | 951 | 183 | 3 |  |  |  |  |  |  |
| 902 | 134 | 3 | 952 | 184 | 3 |  |  |  |  |  |  |
| 903 | 135 | 3 | 953 | 185 | 3 |  |  |  |  |  |  |
| 904 | 136 | 3 | 954 | 186 | 3 |  |  |  |  |  |  |
| 905 | 137 | 3 | 955 | 187 | 3 |  |  |  |  |  |  |
| 906 | 138 | 3 | 956 | 188 | 3 |  |  |  |  |  |  |
| 907 | 139 | 3 | 957 | 189 | 3 |  |  |  |  |  |  |
| 908 | 140 | 3 | 958 | 190 | 3 |  |  |  |  |  |  |
| 909 | 141 | 3 | 959 | 191 | 3 |  |  |  |  |  |  |
| 910 | 142 | 3 | 960 | 192 | 3 |  |  |  |  |  |  |
| 911 | 143 | 3 | 961 | 193 | 3 |  |  |  |  |  |  |
| 912 | 144 | 3 | 962 | 194 | 3 |  |  |  |  |  |  |
| 913 | 145 | 3 | 963 | 195 | 3 |  |  |  |  |  |  |
| 914 | 146 | 3 | 964 | 196 | 3 |  |  |  |  |  |  |
| 915 | 147 | 3 | 965 | 197 | 3 |  |  |  |  |  |  |
| 916 | 148 | 3 | 966 | 198 | 3 |  |  |  |  |  |  |
| 917 | 149 | 3 | 967 | 199 | 3 |  |  |  |  |  |  |
| 918 | 150 | 3 | 968 | 200 | 3 |  |  |  |  |  |  |
| 919 | 151 | 3 | 969 | 201 | 3 |  |  |  |  |  |  |
| 920 | 152 | 3 | 970 | 202 | 3 |  |  |  |  |  |  |
| 921 | 153 | 3 | 971 | 203 | 3 |  |  |  |  |  |  |
| 922 | 154 | 3 | 972 | 204 | 3 |  |  |  |  |  |  |
| 923 | 155 | 3 | 973 | 205 | 3 |  |  |  |  |  |  |
| 924 | 156 | 3 | 974 | 206 | 3 |  |  |  |  |  |  |
| 925 | 157 | 3 | 975 | 207 | 3 |  |  |  |  |  |  |
| 926 | 158 | 3 | 976 | 208 | 3 |  |  |  |  |  |  |
| 927 | 159 | 3 | 977 | 209 | 3 |  |  |  |  |  |  |
| 928 | 160 | 3 | 978 | 210 | 3 |  |  |  |  |  |  |
| 929 | 161 | 3 | 979 | 211 | 3 |  |  |  |  |  |  |
| 930 | 162 | 3 | 980 | 212 | 3 |  |  |  |  |  |  |
| 931 | 163 | 3 | 981 | 213 | 3 |  |  |  |  |  |  |
| 932 | 164 | 3 | 982 | 214 | 3 |  |  |  |  |  |  |
| 933 | 165 | 3 | 983 | 215 | 3 |  |  |  |  |  |  |
| 934 | 166 | 3 | 984 | 216 | 3 |  |  |  |  |  |  |
| 935 | 167 | 3 | 985 | 217 | 3 |  |  |  |  |  |  |
| 936 | 168 | 3 | 986 | 218 | 3 |  |  |  |  |  |  |
| 937 | 169 | 3 | 987 | 219 | 3 |  |  |  |  |  |  |
| 938 | 170 | 3 | 988 | 220 | 3 |  |  |  |  |  |  |
| 939 | 171 | 3 | 989 | 221 | 3 |  |  |  |  |  |  |
| 940 | 172 | 3 | 990 | 222 | 3 |  |  |  |  |  |  |
| 941 | 173 | 3 | 991 | 223 | 3 |  |  |  |  |  |  |
| 942 | 174 | 3 | 992 | 224 | 3 |  |  |  |  |  |  |
| 943 | 175 | 3 | 993 | 225 | 3 |  |  |  |  |  |  |
| 944 | 176 | 3 | 994 | 226 | 3 |  |  |  |  |  |  |
| 945 | 177 | 3 | 995 | 227 | 3 |  |  |  |  |  |  |
| 946 | 178 | 3 | 996 | 228 | 3 |  |  |  |  |  |  |
| 947 | 179 | 3 | 997 | 229 | 3 |  |  |  |  |  |  |
| 948 | 180 | 3 | 998 | 230 | 3 |  |  |  |  |  |  |
| 949 | 181 | 3 | 999 | 231 | 3 |  |  |  |  |  |  |

Send

#46327,aaa					Year 2003 and 2004 – HAVPro Serial Numbers

followed by

#57391,bbb

| SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 03001 | 185 | 11 | 03051 | 235 | 11 | 04001 | 161 | 15 | 04051 | 211 | 15 |
| 03002 | 186 | 11 | 03052 | 236 | 11 | 04002 | 162 | 15 | 04052 | 212 | 15 |
| 03003 | 187 | 11 | 03053 | 237 | 11 | 04003 | 163 | 15 | 04053 | 213 | 15 |
| 03004 | 188 | 11 | 03054 | 238 | 11 | 04004 | 164 | 15 | 04054 | 214 | 15 |
| 03005 | 189 | 11 | 03055 | 239 | 11 | 04005 | 165 | 15 | 04055 | 215 | 15 |
| 03006 | 190 | 11 | 03056 | 240 | 11 | 04006 | 166 | 15 | 04056 | 216 | 15 |
| 03007 | 191 | 11 | 03057 | 241 | 11 | 04007 | 167 | 15 | 04057 | 217 | 15 |
| 03008 | 192 | 11 | 03058 | 242 | 11 | 04008 | 168 | 15 | 04058 | 218 | 15 |
| 03009 | 193 | 11 | 03059 | 243 | 11 | 04009 | 169 | 15 | 04059 | 219 | 15 |
| 03010 | 194 | 11 | 03060 | 244 | 11 | 04010 | 170 | 15 | 04060 | 220 | 15 |
| 03011 | 195 | 11 | 03061 | 245 | 11 | 04011 | 171 | 15 | 04061 | 221 | 15 |
| 03012 | 196 | 11 | 03062 | 246 | 11 | 04012 | 172 | 15 | 04062 | 222 | 15 |
| 03013 | 197 | 11 | 03063 | 247 | 11 | 04013 | 173 | 15 | 04063 | 223 | 15 |
| 03014 | 198 | 11 | 03064 | 248 | 11 | 04014 | 174 | 15 | 04064 | 224 | 15 |
| 03015 | 199 | 11 | 03065 | 249 | 11 | 04015 | 175 | 15 | 04065 | 225 | 15 |
| 03016 | 200 | 11 | 03066 | 250 | 11 | 04016 | 176 | 15 | 04066 | 226 | 15 |
| 03017 | 201 | 11 | 03067 | 251 | 11 | 04017 | 177 | 15 | 04067 | 227 | 15 |
| 03018 | 202 | 11 | 03068 | 252 | 11 | 04018 | 178 | 15 | 04068 | 228 | 15 |
| 03019 | 203 | 11 | 03069 | 253 | 11 | 04019 | 179 | 15 | 04069 | 229 | 15 |
| 03020 | 204 | 11 | 03070 | 254 | 11 | 04020 | 180 | 15 | 04070 | 230 | 15 |
| 03021 | 205 | 11 | 03071 | 255 | 11 | 04021 | 181 | 15 | 04071 | 231 | 15 |
| 03022 | 206 | 11 | 03072 | 0 | 12 | 04022 | 182 | 15 | 04072 | 232 | 15 |
| 03023 | 207 | 11 | 03073 | 1 | 12 | 04023 | 183 | 15 | 04073 | 233 | 15 |
| 03024 | 208 | 11 | 03074 | 2 | 12 | 04024 | 184 | 15 | 04074 | 234 | 15 |
| 03025 | 209 | 11 | 03075 | 3 | 12 | 04025 | 185 | 15 | 04075 | 235 | 15 |
| 03026 | 210 | 11 | 03076 | 4 | 12 | 04026 | 186 | 15 | 04076 | 236 | 15 |
| 03027 | 211 | 11 | 03077 | 5 | 12 | 04027 | 187 | 15 | 04077 | 237 | 15 |
| 03028 | 212 | 11 | 03078 | 6 | 12 | 04028 | 188 | 15 | 04078 | 238 | 15 |
| 03029 | 213 | 11 | 03079 | 7 | 12 | 04029 | 189 | 15 | 04079 | 239 | 15 |
| 03030 | 214 | 11 | 03080 | 8 | 12 | 04030 | 190 | 15 | 04080 | 240 | 15 |
| 03031 | 215 | 11 | 03081 | 9 | 12 | 04031 | 191 | 15 | 04081 | 241 | 15 |
| 03032 | 216 | 11 | 03082 | 10 | 12 | 04032 | 192 | 15 | 04082 | 242 | 15 |
| 03033 | 217 | 11 | 03083 | 11 | 12 | 04033 | 193 | 15 | 04083 | 243 | 15 |
| 03034 | 218 | 11 | 03084 | 12 | 12 | 04034 | 194 | 15 | 04084 | 244 | 15 |
| 03035 | 219 | 11 | 03085 | 13 | 12 | 04035 | 195 | 15 | 04085 | 245 | 15 |
| 03036 | 220 | 11 | 03086 | 14 | 12 | 04036 | 196 | 15 | 04086 | 246 | 15 |
| 03037 | 221 | 11 | 03087 | 15 | 12 | 04037 | 197 | 15 | 04087 | 247 | 15 |
| 03038 | 222 | 11 | 03088 | 16 | 12 | 04038 | 198 | 15 | 04088 | 248 | 15 |
| 03039 | 223 | 11 | 03089 | 17 | 12 | 04039 | 199 | 15 | 04089 | 249 | 15 |
| 03040 | 224 | 11 | 03090 | 18 | 12 | 04040 | 200 | 15 | 04090 | 250 | 15 |
| 03041 | 225 | 11 | 03091 | 19 | 12 | 04041 | 201 | 15 | 04091 | 251 | 15 |
| 03042 | 226 | 11 | 03092 | 20 | 12 | 04042 | 202 | 15 | 04092 | 252 | 15 |
| 03043 | 227 | 11 | 03093 | 21 | 12 | 04043 | 203 | 15 | 04093 | 253 | 15 |
| 03044 | 228 | 11 | 03094 | 22 | 12 | 04044 | 204 | 15 | 04094 | 254 | 15 |
| 03045 | 229 | 11 | 03095 | 23 | 12 | 04045 | 205 | 15 | 04095 | 255 | 15 |
| 03046 | 230 | 11 | 03096 | 24 | 12 | 04046 | 206 | 15 | 04096 | 0 | 16 |
| 03047 | 231 | 11 | 03097 | 25 | 12 | 04047 | 207 | 15 | 04097 | 1 | 16 |
| 03048 | 232 | 11 | 03098 | 26 | 12 | 04048 | 208 | 15 | 04098 | 2 | 16 |
| 03049 | 233 | 11 | 03099 | 27 | 12 | 04049 | 209 | 15 | 04099 | 3 | 16 |
| 03050 | 234 | 11 | 03100 | 28 | 12 | 04050 | 210 | 15 | 04100 | 4 | 16 |

Send

#46327,aaa					Year 2005 and 2006 – HAVPro Serial Numbers

followed by

#57391,bbb

| SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb | SERIAL NUMBER | aaa | bbb |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 05001 | 137 | 19 | 05051 | 187 | 19 | 06001 | 113 | 23 | 06051 | 163 | 23 |
| 05002 | 138 | 19 | 05052 | 188 | 19 | 06002 | 114 | 23 | 06052 | 164 | 23 |
| 05003 | 139 | 19 | 05053 | 189 | 19 | 06003 | 115 | 23 | 06053 | 165 | 23 |
| 05004 | 140 | 19 | 05054 | 190 | 19 | 06004 | 116 | 23 | 06054 | 166 | 23 |
| 05005 | 141 | 19 | 05055 | 191 | 19 | 06005 | 117 | 23 | 06055 | 167 | 23 |
| 05006 | 142 | 19 | 05056 | 192 | 19 | 06006 | 118 | 23 | 06056 | 168 | 23 |
| 05007 | 143 | 19 | 05057 | 193 | 19 | 06007 | 119 | 23 | 06057 | 169 | 23 |
| 05008 | 144 | 19 | 05058 | 194 | 19 | 06008 | 120 | 23 | 06058 | 170 | 23 |
| 05009 | 145 | 19 | 05059 | 195 | 19 | 06009 | 121 | 23 | 06059 | 171 | 23 |
| 05010 | 146 | 19 | 05060 | 196 | 19 | 06010 | 122 | 23 | 06060 | 172 | 23 |
| 05011 | 147 | 19 | 05061 | 197 | 19 | 06011 | 123 | 23 | 06061 | 173 | 23 |
| 05012 | 148 | 19 | 05062 | 198 | 19 | 06012 | 124 | 23 | 06062 | 174 | 23 |
| 05013 | 149 | 19 | 05063 | 199 | 19 | 06013 | 125 | 23 | 06063 | 175 | 23 |
| 05014 | 150 | 19 | 05064 | 200 | 19 | 06014 | 126 | 23 | 06064 | 176 | 23 |
| 05015 | 151 | 19 | 05065 | 201 | 19 | 06015 | 127 | 23 | 06065 | 177 | 23 |
| 05016 | 152 | 19 | 05066 | 202 | 19 | 06016 | 128 | 23 | 06066 | 178 | 23 |
| 05017 | 153 | 19 | 05067 | 203 | 19 | 06017 | 129 | 23 | 06067 | 179 | 23 |
| 05018 | 154 | 19 | 05068 | 204 | 19 | 06018 | 130 | 23 | 06068 | 180 | 23 |
| 05019 | 155 | 19 | 05069 | 205 | 19 | 06019 | 131 | 23 | 06069 | 181 | 23 |
| 05020 | 156 | 19 | 05070 | 206 | 19 | 06020 | 132 | 23 | 06070 | 182 | 23 |
| 05021 | 157 | 19 | 05071 | 207 | 19 | 06021 | 133 | 23 | 06071 | 183 | 23 |
| 05022 | 158 | 19 | 05072 | 208 | 19 | 06022 | 134 | 23 | 06072 | 184 | 23 |
| 05023 | 159 | 19 | 05073 | 209 | 19 | 06023 | 135 | 23 | 06073 | 185 | 23 |
| 05024 | 160 | 19 | 05074 | 210 | 19 | 06024 | 136 | 23 | 06074 | 186 | 23 |
| 05025 | 161 | 19 | 05075 | 211 | 19 | 06025 | 137 | 23 | 06075 | 187 | 23 |
| 05026 | 162 | 19 | 05076 | 212 | 19 | 06026 | 138 | 23 | 06076 | 188 | 23 |
| 05027 | 163 | 19 | 05077 | 213 | 19 | 06027 | 139 | 23 | 06077 | 189 | 23 |
| 05028 | 164 | 19 | 05078 | 214 | 19 | 06028 | 140 | 23 | 06078 | 190 | 23 |
| 05029 | 165 | 19 | 05079 | 215 | 19 | 06029 | 141 | 23 | 06079 | 191 | 23 |
| 05030 | 166 | 19 | 05080 | 216 | 19 | 06030 | 142 | 23 | 06080 | 192 | 23 |
| 05031 | 167 | 19 | 05081 | 217 | 19 | 06031 | 143 | 23 | 06081 | 193 | 23 |
| 05032 | 168 | 19 | 05082 | 218 | 19 | 06032 | 144 | 23 | 06082 | 194 | 23 |
| 05033 | 169 | 19 | 05083 | 219 | 19 | 06033 | 145 | 23 | 06083 | 195 | 23 |
| 05034 | 170 | 19 | 05084 | 220 | 19 | 06034 | 146 | 23 | 06084 | 196 | 23 |
| 05035 | 171 | 19 | 05085 | 221 | 19 | 06035 | 147 | 23 | 06085 | 197 | 23 |
| 05036 | 172 | 19 | 05086 | 222 | 19 | 06036 | 148 | 23 | 06086 | 198 | 23 |
| 05037 | 173 | 19 | 05087 | 223 | 19 | 06037 | 149 | 23 | 06087 | 199 | 23 |
| 05038 | 174 | 19 | 05088 | 224 | 19 | 06038 | 150 | 23 | 06088 | 200 | 23 |
| 05039 | 175 | 19 | 05089 | 225 | 19 | 06039 | 151 | 23 | 06089 | 201 | 23 |
| 05040 | 176 | 19 | 05090 | 226 | 19 | 06040 | 152 | 23 | 06090 | 202 | 23 |
| 05041 | 177 | 19 | 05091 | 227 | 19 | 06041 | 153 | 23 | 06091 | 203 | 23 |
| 05042 | 178 | 19 | 05092 | 228 | 19 | 06042 | 154 | 23 | 06092 | 204 | 23 |
| 05043 | 179 | 19 | 05093 | 229 | 19 | 06043 | 155 | 23 | 06093 | 205 | 23 |
| 05044 | 180 | 19 | 05094 | 230 | 19 | 06044 | 156 | 23 | 06094 | 206 | 23 |
| 05045 | 181 | 19 | 05095 | 231 | 19 | 06045 | 157 | 23 | 06095 | 207 | 23 |
| 05046 | 182 | 19 | 05096 | 232 | 19 | 06046 | 158 | 23 | 06096 | 208 | 23 |
| 05047 | 183 | 19 | 05097 | 233 | 19 | 06047 | 159 | 23 | 06097 | 209 | 23 |
| 05048 | 184 | 19 | 05098 | 234 | 19 | 06048 | 160 | 23 | 06098 | 210 | 23 |
| 05049 | 185 | 19 | 05099 | 235 | 19 | 06049 | 161 | 23 | 06099 | 211 | 23 |
| 05050 | 186 | 19 | 05100 | 236 | 19 | 06050 | 162 | 23 | 06100 | 212 | 23 |

  - Measure the Power Supply Current

- Turn off the DUT.  Connect a CBL121 cable (HVM100 input, Lemo to 3 BNC) to the DUT.

- Use three of the female BNC test adapters shown below to connect a 4990 ohm resister  across each of the CBL121 cable’s male BNC connectors.

BNC Female with isolated ground

Amphenol 31-10

LD Part# 1275.0005

4990 ohm resistor

¼ Watt, 1%

LD Part# 4935.4991

- Setup the bench-top DC power supply with the output voltage set to 10.0V and the current limit adjusted to approximately 500mA.  Turn off the bench-top supply.

- Connect the bench-top supply to the EXTERNAL POWER input of the AHVM100.2.  Make sure the polarity of the connections is correct.

- Verify that the power supply current with an EXTERNAL 10.0V supply is as shown below.  Record the measurements on the HVM100 FINAL CHECKLIST.

  - DUT OFF				 8mA

  - DUT ON with Direct Input		 55 mA

  - DUT ON with ICP Input		 95 mA

- Turn off the DUT and disconnect the bench-top supply.

- Setup the Bench-top DC power supply with the output voltage set to 3.0V.  Turn off the bench-top supply.

- Connect the bench-top supply to the BATTERY TERMINAL inputs of the DUT.  Make sure the polarity of the connections is correct

- Verify that the power supply current with a BATTERY 3.0V supply is as shown below.  Record the measurements on the HVM100 FINAL CHECKLIST.

  - DUT OFF				 100A

  - DUT ON with Direct Input		 100 mA

  - DUT ON with ICP Input		 210 mA

  - Test ICP Current Sources

- Use the DUT to measure the voltage at each ICP input.  The DUT ICP inputs should read between 9.5V and 10.5V.  Record the measurements on the HVM100 FINAL CHECKLIST.

- Turn off the bench-top supply and disconnect it from the DUT.  Disconnect the CBL121 cable from the DUT.

  - Calibrate and Test the DUT Gain Circuitry

- Connect an external power adapter (i.e. PSA027) to the DUT.  Turn on the DUT. Set the DUT baud rate to 9600.

- Connect a CBL106 cable (Direct input, Lemo to BNC) to the DUT and to the test station’s 2209.

- Start the HVMTest.exe software and follow these steps to start the Gain test.

  - Open the Commands menu and select Connect.  The software will report that it has successfully connected to the DUT.

  - Open the File menu and select Customize Tests to open the “Customize Tests” dialog box.

  - Press the Defaults button and then press OK to close the dialog box.

  - Check the Gain Test check box.

  - Open the Commands menu and select Run.

  - The test will run for about 3 minutes and then a “Switch to Charge” prompt will appear.  Replace the CBL106 cable with a CBL107 cable (Charge input, Lemo to BNC) and press Enter to continue the test.  The test will run for about 3 more minutes.

IF ANY TESTS FAIL, TROUBLESHOOT THE PROBLEM BEFORE CONTINUING.

- Disconnect from the HVMTest.exe by opening the Commands menu and selecting Disconnect.   When prompted with “Save New Test Data to Database”, select yes.

  - Test the DUT Charge Amplifiers

- Use the SG.exe software to setup the 2900 function generator to produce a 114.0 dBuV sine wave at 0.4 Hz, see steps below.

  - Start the SG.exe software.

  - Set the SigGen Port to 2 and set the 2209 Port to 0.

  - Check the Connect box to connect to the 2900.  Wait for the Connected message.

  - Enter a frequency of 0.4 Hz and an amplitude of 114.0 dBuV.  Press Send Setup.

- Stop and Reset the DUT.  Set the gain for the X, Y, and Z channels of the DUT to 0 dB.  Set the Display Units to “dB”.  Start the DUT running and wait for the Aeq to run for 10 seconds.  Press Stop.  Verify that the Aeq for the X, Y, and Z channels is between 150.6 and 151.4 dB.  Record the measurements on the HVM100 FINAL CHECKLIST.

- Setup the X, Y, and Z channels of the DUT with 20 dB gain.  Reset the DUT. Start the DUT running and wait for the Aeq to run for 10 seconds.  Press Stop.  Verify that the Aeq for the X, Y, and Z channels is between 150.6 and 151.4 dB. Record the measurements on the HVM100 FINAL CHECKLIST

- Press the Quit button in the SG.exe software.

  - Perform Automated Tests

- Replace the CBL107 cable with a CBL106 cable (Direct input, Lemo to BNC).

- Reconnect to the HVMTest.exe software by opening the Commands menu and selecting Connect.  The software will report that it has successfully connected to the DUT.

- Make sure that only the following check boxes are marked.

LogLin

Crest Test

Burst Test

Detector Test

Fb

- Open the Commands menu and select Run.  The test will run for about 25 minutes.

IF ANY TESTS FAIL, TROUBLESHOOT THE PROBLEM BEFORE CONTINUING.

- Disconnect from the HVMTest.exe by opening the Commands menu and selecting Disconnect.   When prompted with “Save New Test Data to Database”, select yes.

9.0	INSPECTION

No further inspection is required.

10.0	RECORDS

Keep the HVM100 FINAL CHECKLIST (see D0001.8099) on file for the warranty term of the new instrument.

11.0	DISTRIBUTION

Manufacturing

12.0	ATTACHMENTS

Not applicable to this procedure.

13.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
|  | A | 01/05/00 | CBS | Initial Version |
|  | B | 4/19/00 | CBS | Modified for new L1, L2, L3 versions of HVM100 |
| 493 | C | 1/10/03 | CBS | Modified for use with IHVM100 and HAVPro |