HVM100 FINAL CHECKLIST			               	Sales Order

Cal Date		     																						Test Station

Instrument Data, Record the Following Information

Type (√ one)   HVM100	 HAVPro  		SN#:        			Firmware:

Language Option(√ one):	 L1		 L2 		 L3		 L4

Firmware Option (√ one):	 None	 HA (or GT or HAGT)		 WB		 ALL (or HAWB)

Option Codes (enter code value):	#2,     		#3,     		#4,     		#5,

Power Consumption with External = 10.0V

OFF        mA ( 8mA)				ON with Direct Input	  	      mA ( 55 mA)

ON with ICP Input	  	      mA ( 95 mA)

Power Consumption with Vbatt = 3.0V

OFF        A ( 100A)				ON with Direct Input	        mA ( 100 mA)

ON with ICP Input	        mA ( 210 mA)

ICP Current Source Measurements with CBL121 Cable and 4990 ohm Adapters

X         Volts		Y         Volts		Z         Volts	(9.5 to 10.5 Volts)

Charge Amplifiers at 0.4 Hz and 114.0 dBV

20 dB Gain			0 dB Gain

X Aeq		      dB				      dB	(150.6 to 151.4 dB)

Y Aeq		      dB				      dB	(150.6 to 151.4 dB)

Z Aeq		      dB				      dB	(150.6 to 151.4 dB)

Final Inspection:

Inspect Instrument								Package for Shipment

Case screws (5)			     				Instrument

Large rubber foot		     				Manual

Small rubber feet (2)	     				Two AA Batteries

Serial number label		     				Certificate

Calibration label			      				(NOTE: Initial HAVPro calibration is for 15 months)

Battery door				     				I706.07 Warning note if

Camera mount			     				  using Korrvu packaging

Jewel

All Files/Setups Erased

Note: Do not include a Technician Final Inspection slip when checking Havpro’s.

TECHNICIAN				     								DATE

LOGISTICS					     								DATE

QUALITY ASSURANCE	     								DATE