MPR002 TEST PROCEDURE

1.0	PURPOSE AND SCOPE

This document describes the test procedure for verifying that the MPR002 is functioning within specified parameters.  All employees who have responsibility for testing this component are required the follow the instructions detailed in this procedure.

2.0	AFFECTED DEPARTMENTS

Manufacturing

3.0	REFERENCE DOCUMENTS

- Current revision assembly drawing for the MPR002

- SPARK & BLAZE User Manual (part# I706.01)

4.0	RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specifications.

- Troubleshoot and correct product as required.

- Communicate concerns to the Supervisor of Quality Assurance.

- Request management review of product concerns.

- Follow established ESD standards.

5.0	DEFINITIONS

MPR002: Spark 3” preamplifier tube with attached MIC002 microphone.

6.0	SAFETY PRECAUTIONS

Safety glasses when soldering, lead clipping, or testing power supplies.

7.0	EQUIPMENT AND MATERIALS

## Larson Davis Spark Dosimeter (703, 703+, 704, 705, 705+, 706 or 706RC)

## ADP018 Adapter

- CBL066 Cable (BNC to BNC)

- JetEye PC IrDA Serial Communications Adapter (DVX008)

- Computer (PC 486 66MHz or newer, running Windows 95 or newer)

- SLMTestProd.exe Windows software (Part #5499.0004 – current version)

- Larson Davis Test Station (2900, 2209, Computer)

8.0	INSTRUCTIONS

  - Test the MPR002 with the SLM Test Software

- Remove the MIC002 and connect an ADP018 to the MPR002 being tested.

- Attach the MPR002 to a previously tested Spark instrument.

- Connect the CBL066 cable from the ADP018 to the Larson Davis Test Station.

- Place the JetEye in front of the Spark dosimeter’s IR window.  Turn on the Spark.

- Turn on the PC’s IR port (see Spark and Blaze User manual for details).

- Start the SLMTestProd.exe software and follow these steps to start the test.

  - Use the “Instruments” selection box to select Spark Family.

  - Open the Commands menu and select Connect.  The software will report that it has successfully connected to the SPARK.

  - Open the File menu and select Customize Tests to open “Customize Tests” dialog box.

  - Press the Defaults button to set the test parameters.

  - Verify that the following settings are as follows and press OK to close the dialog box.

    - Weight Test Level:		130

    - Gain Level for Tests		0

    - Log Lin Weight:		A Weight

    - Log Lin Detector:		Slow

    - Log Lin Cal Frequency:	1000 Hz

    - Log Lin Test Frequency:	1000 Hz

    - Log Lin Step Incr.:		3

    - Calibrate:			Not Checked

    - No Calibration:		Checked

    - Ask To Calibrate:		Not Checked

    - Do Time Average Test:	Not Checked

  - Make sure that only the  Loglin and C Weight tests are selected.

  - Verify that Do Scale-offset is selected in the Instrument Adjustments box.

  - Open the Commands menu and select Run.

- Verify that both tests passed.  Disconnect from the SLMTestProd.exe by opening the Commands menu and selecting Disconnect.   When prompted with “Save New Test Data to Database”, select NO.

- Firmly tighten the MIC002 microphone onto the preamp using some soft rubber sheeting or a similar tool.

9.0	INSPECTION

No further inspection of the MPR002 is required.

10.0	RECORDS

Since the MPR002 is not serialized, no test records are stored.

11.0	DISTRIBUTION

Manufacturing

12.0	ATTACHMENTS

Not applicable to this procedure.

13.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
|  | A | 9/27/00 | CBS | Initial Version |