PRA950 FAMILY FINAL TEST

- PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to ensure to test and verify the function of the members of the PRA950 Family.

- AFFECTED DEPARTMENTS

Manufacturing

3.0	REFERENCE DOCUMENTS

PRA950 schematics for the version & revision being tested.

4.0	RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specification.

- Troubleshoot and correct product as required.

- Communicate concerns to Supervisor or Quality Assurance.

- Request management review of product concerns.

- Use proper ESD protective equipment.

5.0	DEFINITIONS

ESD -	Electro-Static Discharge

6.0	SAFETY PRECAUTIONS

- Safety Glasses are required when soldering, lead clipping, or testing power supplies.

- Follow general electrical precautions for working with energized, low voltage circuits

7.0	EQUIPMENT & MATERIALS

Frequency test fixture

Current test fixture

Trim resistor fixture

Function generator

Oscilloscope

2900

2200

Malco connector cable

Breakout box

Extension cable

8.0	INSPECTION

- Examine assembly for misloaded parts, shorts, etc.

- Check PRA950-2 and PRA950–4 units and verify that a jumper wire has been loaded in place of the CR1 and CR2 components.

9.0	INSTRUCTIONS

If the testing is being performed on an instrument that has been returned for re-certification or repair, skip to section 9.2.

There are three fixtures used in the instructions, each is described and/or illustrated in the following text.

Current Test Fixture:

BNC

Malco Connector with Cable

1 Kohm

Trim Resistor Fixture:

20 Kohm Pot

Frequency Test Fixture:

BNC

Malco Connector with Cable

4.02 Kohm

10 F Cap

### 9.1 Resistor trim

- Remove the Pre-amp body from the unit to be trimmed and assemble it back together without the body, using a jumper cable in place of the body.

- Connect the extension cable to the 2200 for (PRA950-2 and –4 units) and to the unit under test. Use a 2900 for (PRA950-L2 and L-4 units).

- Connect the Malco test cable to the unit under test and connect the Current test fixture to the other end of the malco cable.

- Connect the Trim resistor fixture onto the unit under test at R1.

- Turn on either the 2200 or the 2900; connect your voltmeter probes across the resistor on the Current test fixture. Adjust the trim pot on the resistor fixture until you have one of the following readings:

PRA950-2    = 	2.000 +/- .05 Vdc

PRA950-4    = 	4.000 +/- .10 Vdc

### PRA950-L2  = 	2.000 +/- .05 Vdc

PRA950-L4  = 	4.000 +/- .10 Vdc

- Turn off the 2200 or 2900 and remove the current fixture and measure the resistance across the trim pot on the resistor fixture.

- Select an equivalent resistor size and solder it into R1

- Assemble the unit under test using the Pre-amp body and check the voltage reading. Record the voltage value on the "PRA950 Family Final Checklist" (D0001.8158-1).  If within tolerance levels continue if not select a closer resistor value.

9.2	AC SIGNAL DC VOLTAGE TEST for New and Service

- With the Current test fixture connected, measure the DC voltage and note it on the final Checklist.

- Connect the Frequency test fixture.

- Apply a 1 Volt(rms) at 1 KHz signal from the generator and into the capacitor on the Frequency fixture.

- View the signal out, the level should be 0.95 - 1.00 Vac.  Record the voltage on the checklist.

9.3	Finish

- Testing is completed.  Turn off the power to the unit and remove the unit under test.

- If unit passes mark the "Pass" box on the checklist.  If unit fails mark the "Fail" box.

- Verify that the model, serial number and date have been entered onto the checklist.

- For passing units, generate a certificate of compliance

- Technician signs the checklist and the certificate

- Certificate goes to customer, checklist is filed by the technician.

- RECORDS

- Complete and turn in the technician inspection and rework report.

- Complete checklist and file appropriately.  The technician group will maintain checklist for the life of the instrument warranty or re-certification interval.

11.0	DISTRIBUTION

Manufacturing

12.0	ATTACHMENTS

None

13.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 361 | A | 11/14/01 | RH | Initial Release |
| 1143A | B | 9/29/08 | EO | Updating procedure to clarify procedure |