PRA951 FAMILY FINAL TEST CHECKLIST

Model:

Serial:

Date:

Refer to work instruction D0001.8163 for detailed information on each test to be performed.

New Instrument				Re-certification

In Tolerance as received?	  Yes	  No

Resistor Trim

Adjust the trim pot on the fixture to the reading that matches the model being tested and record the voltage.

Voltage

PRA951-2	=	2.000 +/-  .05 Vdc

PRA951-4	=	4.000  +/- .10 Vdc

PRA951-L2	=	2.000  +/- .05 Vdc

PRA951-L4	=	4.000  +/- .10 Vdc

Signal Voltage Test

Measured Voltage:

Unit Status:

Pass					Fail

Technician: