BAT007

- PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to ensure that the BAT007 charges and holds a charge.

- AFFECTED DEPARTMENTS

Manufacturing

3.0	REFERENCE DOCUMENTS

None

4.0	RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specification.

- Troubleshoot and correct product as required.

- Communicate concerns to Supervisor or Quality Assurance.

- Request management review of product concerns.

Use proper ESD protective equipment.

5.0	DEFINITIONS

- BOLD words represent a hard key function (the keys permanently labeled on the keyboard).

- BOLD[A] words represent a soft key function (the BOLD represents the text on the screen and [A] identifies the hard key associated with the function).

6.0	SAFETY PRECAUTIONS

No process specific safety precautions needed.

7.0	EQUIPMENT & MATERIALS

- PSA013

- PSA004

- 2800, 2900, or 3000+

8.0	INSPECTION

The process is an inspection of the BAT007.  No further inspection beyond the requirements in the instructions is required.

9.0	INSTRUCTIONS

- Charge the battery by plugging the PSA004 into the PAS013, and then placing the battery into the PSA013.

- The red LED on the PSA013 should come on.

NOTE:  On some batteries that are new and being charged for the first time, the charger will turn off after a few minutes.  If this happens, remove the battery then re-insert it.

- It should take about 4½ to 5½ hours to charge.

- Put the battery into a 2800, 2900, or 3000+.

- Turn on the unit, turn on the backlight, and run the unit.

Keys:  ON, -, SHIFT, SCREEN, BKLT ON[C], EXIT, R/S

- Press SYSTEM, EXIT, SYSTEM

- The 2900 should read a voltage of about 7.8 V.

- Turn off the 2900.

Keys:  R/S, SHIFT, OFF

- Remove the battery from the 2900.

- Place a “BAT007” label on the bottom of the battery.  If you have to make some they are in the ‘Options’ label located in the ‘3200lab’ directory in the ‘My Label’ program.

- RECORDS

No records are generated for this process.

11.0	DISTRIBUTION

Manufacturing

Engineering

12.0	ATTACHMENTS

None

13.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 304 | A | 6/15/01 | JG | Initial Release |