Mic Lab Humidity Chamber Maintenance

Manufacturer:	Espec

Model:		LHU-112

Location:		Microphone Lab

Pre-operation Maintenance: Check water level in water tank and refill as needed.

Weekly Steps:

- Clean wick pan and check the water level.

- Clean humidifying tray.

- Check wet bulb wick. If soiled replace.

Monthly Steps – after performing steps one through three, do the following:

- Clean the water circuit. Change the water in the tank, and clean the inside of the tank. Clean inside the humidifying tray water leveler with a clean cloth.

- Clean the air intake and exhaust port on the rear panel.

- Clean the condenser fins.

- Clean the electromagnet pump filter.

- Check function of the automatic overheat protective circuit.

- Check function of the leak breaker.