CAL200 Service Checklist

| Serial Number: | Date: |
| --- | --- |
| Customer: |  |
| Calibrator Class | Technician: |

| Adapters: | ¼” | 3/8” | None |
| --- | --- | --- | --- |

Check tone

| O-rings: | Ok | Replaced |
| --- | --- | --- |

Incoming levels at 1013 mbar:

| Frequency: 1000 Hz  10 Hz (Class 1L:  30 Hz) |  | In Tolerance |  |
| --- | --- | --- | --- |
| SPL: 114 dB  0.2 dB (Class 1L:  0.3 dB) |  | Out of Tolerance |  |
| SPL 94 dB  0.2 dB (Class 1L:  0.3 dB) |  |  |  |

Pressure test:  Class 1 SPL   0.3 dB 650 to 1080 mbar. / Class 1L print correction

Replace battery

| Final SPL: 114 dB  .04 dB @ 1013 mbar |  |
| --- | --- |
| Final SPL: 94 dB  .04 dB @ 1013 mbar |  |
| Final Frequency: 1000.0  .5 Hz |  |

Distortion test  (Class 1  2% Harmonic 34 dB down or more)

(Class 1L  3% Harmonic 30 dB down or more)

Enter in LD database

Print Certificates

Calibration label on unit

Cal Seal on unit

Notes: