CAL250 Service Checklist

| Serial Number: | Date: |
| --- | --- |
| Customer: |  |
| Technician: |  |

| Adapters: | ¼” | 3/8” | ½” | None |
| --- | --- | --- | --- | --- |

Check tone

Check screen

| O-rings: | Ok | Replaced |
| --- | --- | --- |

Incoming levels @ 1013 mbar:

| Frequency: 251.19 Hz  2 Hz |  | In Tolerance |  |
| --- | --- | --- | --- |
| SPL: 114 dB  .2 dB |  | Out of Tolerance |  |

Pressure Test  650 to 1080 mbar  0.3 dB

Replace battery

LDL calibration seal placed over pot adjustments

| Final SPL @ 1013 mbar: 114 dB  .03 dB |  |
| --- | --- |
| Final Freq @ 1013 mbar: 251.2 Hz  .5 Hz |  |

Check distortion, < 2% (harmonic 34 dB down or more)

Enter in LD database

Print Certificates

Calibration label on unit

Cal Seal on unit

Notes: