CAL250 RE-CERTIFICATION

Receive CAL250

Check tone, screen, and O-ring

Open CAL250 Service Checklist (D0001.8283)

Ok?

Determine problem

Correct Problem

Receive CAL250

Major

Minor

Replace battery

Check incoming level and record on Checklist.  Reference CAL250 Final Certification (D0001.8192) section 4.2 to obtain levels.

Perform pressure test.  Reference CAL250.F Assembly and Test section 4.5.

Needs adjustment?

Print out Certificate of Calibration

Refer to D0001.8192 CAL250 Final Certification Section 4.5

Enter service record in LD database

Generate necessary labels, Print Return to Customer Form

No

Yes

Adjust level

(See diagrams below)

Check levels again

Return equipment & paperwork to service desk

Yes

No

Passes?

Yes

No

Determine problem

Correct problem

Minor

Major

Send quote

Customer accepts

Yes

No

Return unrepaired equipment

Send quote

Older CAL250’s require disassembly to adjust the frequency.  (Place the CAL250 over the microphone and preamp with the circuit board hanging outside the tube so the adjustment potentiometer may be accessed.  Carefully connect a 9V battery to the battery clip, turn the unit on, monitor the frequency on the oscilloscope and adjust to specifications.  Remove the CAL250 from the microphone and preamp and reassemble.  See diagram below:)

End view of tube, rear battery door removed

Amplitude R20			Frequency R19 (Accessible on newer models)

8.0		REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 736 | A | 3/4/04 | SM | Initial Release |
| 1104 | B | 6/16/08 | SM | Updated |
| 1489 | C | 11/13/15 | SM / JH | MET/CAL System |