PRM915 FINAL TEST

- PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to ensure the compliance of the specified product to established specifications.

- AFFECTED DEPARTMENTS

Manufacturing

3.0		REFERENCE DOCUMENTS

Assembly drawings and schematics for the PRM915

4.0		RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specification.

- Troubleshoot and correct product as required.

- Communicate concerns to Supervisor or Quality Assurance.

- Request management review of product concerns.

- Use proper ESD protective equipment.

5.0		DEFINITIONS

None

6.0		SAFETY PRECAUTIONS

- Safety Glasses are required when soldering, lead clipping, or testing power supplies.

- Follow general electrical precautions for working with energized, low voltage circuits

7.0		EQUIPMENT & MATERIALS

# ADP040 (7pin Lemo with signal out)

- ADP041 (7pin Lemo with PRM915 signal input)

- BNC to BNC cable

- CBL071

- ADP005

- 2900 TEST STATION

- INSPECTION

- Examine the cosmetic appearance of the preamplifier

- Verify that the contacts and connectors are properly aligned on the part

- Questions about appearance or quality are to be brought to the attention of the supervisor, manager or QA

- INSTRUCTIONS

9.1		Trim

Refer to the PRM902 PRETEST AND TRIM procedure #D0001.8125.

NOTE: that if the current is wrong you will need to adjust the value of R7.

Also check that the small trace connecting a Via to the trace between L3 and Pin 7 of the Lemo connector is cut as shown below.

9.2		Test #1

Connect to 2900 test station as shown below.

- Run the PreampTest program.

- Select your name in the Technician field.

- For Model: select 915.

- When it asks for the serial number, enter it and press OK.

- Click on the play button.

- When it asks for the serial number, enter it and press OK.

9.3		Test #2

- Connect to 2900 test station as shown below.

- Run the “PreampTest” program

- Select your name in the Technician field.

- For Model: 915 Inject.

- Click on the play button.

- When it asks for the serial number, enter it and press OK.

- Print  both test results when finished

9.4		Test with mic.

- Connect the preamp to the test station as shown in the first diagram, but place a 2540 microphone on the preamp in place of the ADP005.

- Put a CAL250 over the microphone and activate the calibrator.

- The signal should be  120 mV AC.

(This level is dependent on the microphone used it is best to note this level and always use the same microphone for this test)

9.5  	Cert.

- Create & Print Certificate of Conformance for PRM915 (without due date) in the LD database

- Sign cert and include above test results

- Print cal sticker and barcode sticker with s/n and place on outside of blue tube.

- RECORDS

Test data, certificates and other associated records are maintained per the Quality Records Matrix, D0001.1126-1.

11.0	DISTRIBUTION

Available online in work instruction area of Document Control.

12.0	ATTACHMENTS

None

13.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 758 | A | 3/22/04 | JG | Initial Release |
| 1154A | B | 11/11/08 | SC | Clarification of trim procedure |
| 1867 | C | 1/16/19 | SRC | Added clarity so it was more production friendly |