PRM916 FINAL TEST

- PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to ensure the compliance of the specified product to established specifications.

- AFFECTED DEPARTMENTS

Manufacturing

3.0		REFERENCE DOCUMENTS

Applicable assembly drawings and schematics for the PRM916

4.0		RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specification.

- Troubleshoot and correct product as required.

- Communicate concerns to Supervisor or Quality Assurance.

- Request management review of product concerns.

- Use proper ESD protective equipment.

5.0		DEFINITIONS

None

6.0		SAFETY PRECAUTIONS

- Safety Glasses are required when soldering, lead clipping, or testing power supplies.

- Follow general electrical precautions for working with energized, low voltage circuits

7.0		EQUIPMENT & MATERIALS

# ADP040 (7pin Lemo with signal out)

- ADP041 (7pin Lemo with PRM915 signal input)

- TWO BNC to BNC cable’s (one must be very short)

- CBL071

- PRM916 SHIELD (A shorted ADP005 with No Capacitor)

- ADP005

- 2900 Test station with the Preamp Test software installed.

- INSPECTION

- Examine the cosmetic appearance of the preamplifier

- Verify that the contacts and connectors are properly aligned on the part

- Questions about appearance or quality are to be brought to the attention of the supervisor, manager or QA

- INSTRUCTIONS

9.1		Trim

Refer to the PRM902 PRETEST AND TRIM procedure #D0001.8125.

NOTE: If the current is wrong you will need to adjust the value of R7.

Also check that the small trace connecting a Via to the trace between L3 and Pin 7 of the Lemo connector is cut as shown below.

9.2		Test of frequency response of the preamp with the signal through ADP005.

- Connect the preamp to the 2900 test station as shown below.

- Run the Larson Davis Preamp Test program.

- Select PRM916.

- Click on the START button.

- When it asks for the serial number, enter it and press OK.

9.3	Test of frequency response of preamp with the signal driving the internal         capacitance bridge.

- Connect to 2900 test station as shown below.

- Run the Larson Davis Preamp Test program.

- Select PRM916 Injected.

- Click on the START button.

- When it asks for the serial number, enter it and press OK.

- Print  both test results when finished

9.4		Cert.

- Create & Print Certificate of Conformance for PRM916 (without due date) in the LD database

- Sign cert and include above test results

- Print cal sticker and barcode sticker with s/n and place on outside of blue tube.

- RECORDS

Records that are generated as a result of this instruction are maintained per the Quality Records Matrix, D0001.1126-1.

11.0	DISTRIBUTION

Instruction is available in the work instructions area of Document Control.

12.0	ATTACHMENTS

None

13.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 757 | A | 3/22/04 | JG | Initial Release |
| 1155A | B | 11/11/08 | SC | Clarification of trim procedure |
| 1541 | C | 2/17/15 | JG | Changed picture in test 1 to shorten cable. Changed from Lab to the Preamp Test software. Changed heading of sections 9.2 and 9.3. |
| 1867 | D | 1/16/19 | SRC | Added clarity so it was more production friendly |