PCB Piezotronics Inc., Larson Davis Division

Oscilloscope Test & Verification Checklist

| Date: |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Instrument Under Test |  |  |  |  |  |  |  |  |  |
| Manufacturer: |  | Tektronix |  |  |  |  |  |  |  |
| Model: |  | 2236 |  |  | Serial #: |  |  |  |  |
| Calibration Interval: |  |  |  |  |  |  |  |  |  |
| Received Status: | In Tolerance |  |  | Out of Tolerance |  |  |  |  |  |
| Temperature: |  |  |  | Humidity: |  |  |  |  |  |
| TEST INSTRUMENT(S) |  |  |  |  |  |  |  |  |  |
| Manufacturer | Model |  |  | Serial # |  | Traceability |  |  | Calibration Due |
| OSCILLOSCOPE TEST RESULTS |  |  |  |  |  |  |  |  |  |
| Frequency | Voltage |  | Allowed Range |  |  | Measured |  |  | In Tolerance? |
| 1 KHz | 3.162 Volts |  | 3.080 to 3.240 V |  |  |  |  |  |  |
| 1 KHz | 1 Volt |  | 0.974 to 1.026 V |  |  |  |  |  |  |
| 1 KHz | 100 mV |  | 0.0974 to 0.1026 V |  |  |  |  |  |  |
| 20 Hz | 1 Volt |  | 0.974 to 1.026 V |  |  |  |  |  |  |
| 200 Hz | 1 Volt |  | 0.974 to 1.026 V |  |  |  |  |  |  |
| 2 KHz | 1 Volt |  | 0.974 to 1.026 V |  |  |  |  |  |  |
| 20 KHz | 1 Volt |  | 0.974 to 1.026 V |  |  |  |  |  |  |
| MULTIMETER TEST RESULTS Reference Meter; Allowed Range; Measured; In Tolerance?. 1 Volt; 0.974 V to 1.026 V. 0.01 Volt; 9.7 mV to 10.3 mV. 1KOhm; 970 Ohms to 1030 Ohms. Instrument has been tested and verified according to the test procedure, D0001.8318. This instrument has been found to be: In Tolerance Out of Tolerance (remove from service until repaired & re-tested) If In Tolerance, this instrument will be due for re-certification on: |

(Technician)						(Date)