PCB Piezotronics Inc., Larson Davis Division

Handheld DMM Test & Verification Checklist

| Date: |  | FORMTEXT |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Instrument Under Test |  |  |  |  |  |  |  |  |  |  |  |
| Manufacturer: |  |  |  |  |  |  |  |  |  |  |  |
| Model: |  |  |  |  |  | Serial #: |  |  |  |  |  |
| Calibration Interval: |  |  |  |  |  |  |  |  |  |  |  |
| Received Status: |  | In Tolerance |  |  | Out of Tolerance |  |  |  |  |  |  |
| Temperature: |  |  |  |  | Humidity: |  |  |  |  |  |  |
| TEST INSTRUMENT(S) |  |  |  |  |  |  |  |  |  |  |  |
| Manufacturer |  | Model |  |  | Serial # |  | Traceability |  |  | Calibration Due |  |
| HANDHELD DMM TEST RESULTS |  |  |  |  |  |  |  |  |  |  |  |
| Reference Meter | Fluke 87 DMM Allowed Range |  |  | Fluke 21 DMM Allowed Range |  |  |  | Measured |  |  | In Tolerance? |
| 1 Volt | 0.99 V to 1.003 V |  |  | 0.991 V to 1.009 V |  |  |  |  |  |  |  |
| 0.01 Volt | 9.84 mV to 10.16 mV |  |  | 9.78 mV to 10.22 mV |  |  |  |  |  |  |  |
| 1 KOhm | 996 Ohms to 1004 Ohms |  |  | 992 Ohms to 1008 Ohms |  |  |  |  |  |  |  |
| Instrument has been tested and verified according to the test procedure, D0001.8318. This instrument has been found to be: In Tolerance Out of Tolerance (remove from service until repaired & re-tested) If In Tolerance, this instrument will be due for re-certification on: |  |  |  |  |  |  |  |  |  |  |  |

(Technician)						(Date)