NMS017 Final Assembly and Test

1.0		PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to perform the final assembly and test of the NMS017 System.  This process also serves as a verification that the product continues to meet the specifications and criteria established by Engineering.

2.0		AFFECTED DEPARTMENTS

Manufacturing

3.0		REFERENCE DOCUMENTS

- 831 Final Checklist

- EPS03X Assembly drawing

- D0001.8310 831 Final Assy and Test

- D0001.8322	831-Int Final Assy

- D0001.8307  Configure 831/LXT

- D0001.8327  S2102.2

- I2150.02

- I2150.05

- I2151.01

- I2151.02

- PSA012-XX

4.0		RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specification.

- Troubleshoot and correct product as required.

- Communicate concerns to Supervisor or Quality Assurance.

- Request management review of product concerns.

- Use proper ESD protective equipment.

5.0		SAFETY PRECAUTIONS

- Safety Glasses are required when soldering, lead clipping, or testing power supplies.

- Follow general electrical precautions for working with energized, low voltage circuits

6.0		EQUIPMENT & MATERIALS

EPS032  		SST enclosure

CBL160		Cable, Power Solar

377B20		Mic ½”

831			SLM

CBL153		Cable 426A12 to 831-INT control

CBL152		Cable 426A12 to 831 Signal

426A12 		Outdoor Preamp

TRP019		Hinged 17’ pole

426A12-NPT	Coupler

831-INT		831 Interface Unit

831-ELA		Opt 831 Event

831-LOG		Option, 831 data logging

Phillips screw driver

7.0		INSPECTION

Inspect the parts for nonconformities as they are assembled together and subsequently tested.

- INSTRUCTIONS

- Verify firmware options are loaded per D0001.8307.

- Install components as shown in assembly drawing EPS03X.  Bundle cables neatly with Tie wraps where appropriate.  Reference drawing PSA012-XX for solar panel mounting.

- Verify functionality by connecting:

  - CBL153 from 426A12 to 831-INT control.

  - CBL152 from 426A12 to 831 Signal.

  - CBL160 from Solar charger (Label: Batt + and -) to 831-INT (Label: Battery).

Options:

  - Depending on the particular sales order, connect cabling from solar charger (Label: Batt + and -) to battery(s): CBL162 for 2x 21Ah or CBL161 for 100AH underground battery.  Attach temperature sensor to side of battery with tape.

  - Depending on the particular sales order, connect cabling from solar panel (Label + and -) to solar charger (Label: Array + and -) : PSA012-50 for 50W panel or PSA012-80 for 80W panel.

  - For GPS:  Connect GPS USB cable to any 831-INT port labeled “USB Port.”

  - For Modem:  Connect Modem USB cable to any 831-INT port labeled “USB Port.”

  - For Wind Monitor:  Connect Wind monitor cable to 831-INT port labeled “Wind Sensor.”  Reference I2151.01 and I2151.02 for mounting.

  - For Temperature and Humidity Sensor:  Connect T & H Sensor cable to 831-INT port labeled “Sensors A & B.”  Reference I2151.01 and I2151.02 for mounting.

- Disassembly:

- Where appropriate, ship parts separately.

- Field Installation:

1.  Reference field installation drawings I2150.02 and I2150.05.

10.0	DISTRIBUTION

This instruction is available to employees electronically via the online Document Control area.

11.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
|  | A | 9/11/07 | CT | Initial Release |
| 3669 | A | 8/12/08 | AR | ECO RELEASE |