NMS018 Final Assembly and Test

1.0		PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to perform the final assembly and test of the NMS018 System.  This process also serves as a verification that the product continues to meet the specifications and criteria established by Engineering.

2.0		AFFECTED DEPARTMENTS

Manufacturing

3.0		REFERENCE DOCUMENTS

- 831 Final Checklist

- EPS0XX Assembly drawing

- D0001.8310 831 Final Assy and Test

- D0001.8322	831-Int Final Assy

- D0001.8307  Configure 831/LXT

- D0001.8327  S2102.2

- TRP020-XX

- NMS01X

4.0		RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specification.

- Troubleshoot and correct product as required.

- Communicate concerns to Supervisor or Quality Assurance.

- Request management review of product concerns.

- Use proper ESD protective equipment.

5.0		SAFETY PRECAUTIONS

- Safety Glasses are required when soldering, lead clipping, or testing power supplies.

- Follow general electrical precautions for working with energized, low voltage circuits

6.0		EQUIPMENT & MATERIALS

EPS031  		Enclosure

377B20		Mic ½”

831			SLM

CBL153		Cable 426A12 to 831-INT control

CBL152		Cable 426A12 to 831 Signal

CBL147		Cable Battery

426A12 		Outdoor Preamp

TRP019		Hinged 17’ pole

426A12-NPT	Coupler

831-INT		831 Interface Unit

831-ELA		Opt 831 Event

831-LOG		Option, 831 data logging

Phillips screw driver

2 Pipe wrench with cloth(to prevent damaging the surface of the pipe)

7.0		INSPECTION

Inspect the parts for nonconformities as they are assembled together and subsequently tested.

- INSTRUCTIONS

- Verify firmware options are loaded per D0001.8307.

- Drill holes and install components as shown in assembly drawing EPS03X.  Bundle cables neatly with Tie wraps where appropriate.

- Connect Pipe as shown in drawings TRP020-XX and NMS01X.  Use Pipe wrenches with rags/cloth to prevent marring.

- Verify functionality by connecting:

  - CBL153 from 426A12 to 831-INT control.

  - CBL152 from 426A12 to 831 Signal.

  - CBL147 from battery to 831-INT port labeled “Battery.”

  - 0195.00054 Power Supply from 831-INT to Power Strip

  - CBL159 from 831 Auxiliary port to 831-INT port labeled “USB Hub”

Options:

  - For GPS:  Connect GPS USB cable to any 831-INT port labeled “USB Port.”

  - For Modem:  Connect Modem USB cable to any 831-INT port labeled “USB Port.”

  - For Wind Monitor:  Connect Wind monitor cable to 831-INT port labeled “Wind Sensor.”  Reference I2151.01 and I2151.02 for mounting.

  - For Temperature and Humidity Sensor:  Connect T & H Sensor cable to 831-INT port labeled “Sensors A & B.”  Reference I2151.01 and I2151.02 for mounting.

  - For battery back up systems, connect cabling from 831-INT port labeled “battery” to EPS038/9

- Disassembly:

- Where appropriate, ship parts separately.

10.0	DISTRIBUTION

This instruction is available to employees electronically via the online Document Control area.

11.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
|  | A | 9/11/07 | CT | Initial Release |
| 3699 | A | 8/12/08 | AR | ECO RELEASE |