CER-426A12

Outdoor Preamp Environmental Test

1.0	PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to ensure the compliance of the 426A12 to established specifications for temperature and humidity. The procedure is written to apply to any technician that is given the responsibility and authority to test and / or certify this product.

2.0	AFFECTED DEPARTMENTS

Manufacturing

3.0	REFERENCE DOCUMENTS

Operation and Specifications as found in I426A12.01 426A12 Reference Manual

Operating manual for all test equipment used.

A2102.1, A2102.2, A2102.3, S2102.1 and S2102.2 Schematic and Assembly Drawings may be needed.

D0002.0007-1 426A12 Certificate of Environmental Test

4.0	RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specifications.

- Troubleshoot and correct product as required.

- Communicate concerns to the Supervisor of Quality Assurance.

- Request management review of product concerns.

- Follow established ESD standards.

5.0	DEFINITIONS

None applicable to this procedure.

6.0	SAFETY PRECAUTIONS

- This procedure will expose the equipment under test to extreme cold and hot conditions. Use caution and wear gloves when handling equipment that may be very cold or hot to avoid burns.

- Safety glasses are required when soldering, lead clipping, or testing power supplies.

- EQUIPMENT AND MATERIALS

## The “unit under test” is one or more 426A12 and one Model 831 for each 426A12

- The microphone that goes with the 426A12 should be on the 426A12

- One USB cable for each unit under test

- One CBL152 Signal Cable for each 426A12

- One CBL154 Control Cable for each 426A12

- 12 Volt DC power source for 426A12 - SLA battery with 2.5mm power plugs

- USB Hub(s) with one port for each unit under test

- Environmental Test Chamber

- Computer with the chamber program, 426A12.seq sequence file and the ReportGen program

8.0	INSPECTION

Equipment to be environmentally tested should already have been bench certified and thus needs no further inspection at this point.

9.0	INSTRUCTIONS

The purpose of this test is to ensure that the 426A12 will meet the requirements of the specifications for temperature and humidity.  The limits are programmed into the ReportGen program and listed in the below.

426A12 Relative SPL vs. Temperature Limits

Tested Temperature Range:  -40 °C to +70 °C

| Temperature (°C) | Lower Limit (dB) | Upper Limit (dB) | Uncertainty (dB) |
| --- | --- | --- | --- |
| -40 to -20 | -1.1 | +1.1 | 0.1 |
| -10 to 50 | -0.6 | +0.6 | 0.1 |
| 60 to 70 | -1.1 | +1.1 | 0.1 |

426A12 Relative SPL vs. Humidity Limits

Tested Relative Humidity Range:  50% to 96%

±0.6 dB with 0.1 dB uncertainty

426A12 Relative SPL vs. Humidity Endurance Limits

Tested Relative Humidity at 96% and 50 °C for 2 hours

±0.6 dB with 0.1 dB uncertainty

426A12 Relative SPL vs. Humidity Stress Limits

Tested with Relative Humidity at 96% and Temperature Cycled 50 °C and 25 °C

±0.6 dB with 0.1 dB uncertainty

426A12 Internal Humidity vs. Temperature and Humidity Limits

< 45% RH with 5% uncertainty

426A12 Internal Temperature vs. Chamber Temperature Limits

±7 °C with 3 °C uncertainty

9.1	Connect the units for testing.

- Place the 426A12 preamplifiers with microphones attached inside the environmental test chamber.

- Connect the 426A12 preamplifier to the Model 831 with the CBL152.

- Connect the CBL154 control cable for the 426A12 preamplifiers, to the Model 831 and to the 12 Volt battery power via the 2.5mm power jack

- Connect a USB cable from each Model 831 to a USB Hub. Do not daisy-chain hubs, each hub must connect directly to the computer for them to work properly.  Each hub must be a “powered hub”

9.2	Environmental Chamber operation

- Start the Chamber test program “LD Environmental Chamber Control Software” by invoking chamber.exe.

- Click on 426A12 Test tab

- Enter the number of 426A12 preamps attached.

- The 426A12 serial numbers can be detected automatically.  Click on the AutoDetect button.

- Enter the Microphone model numbers and serial numbers (i.e. 377B02/108345).

- Select the 426A12.seq sequence file.

- Click Begin 426A12 test.

9.3	Once the test is completed, Plot the Test Results

- Start the Report Generator program

.

- Select Environmental Test Results as the test type.

- Select the Instrument field to be 426A12.

- Select Serial Number and enter the serial number in the adjacent field and press Lookup or drop down the field and choose the serial number of the unit to generate reports for.

- Locate and view the test results.  Verify that the “uncertainty flags” are within the limit lines.  Any blue flags above or below the limit lines represents a failed condition.

- Print the test documentation to be provided to the user by pressing the  button and selecting the reports from the list provided.  The 426A12 Relative Internal Humidity vs. Chamber Humidity report is not printed.

9.4	Completion of the Environmental Chamber Test

- When the reports indicate the units pass, then the equipment can be removed from the chamber.

- Make a new Certificate of Environmental Test using D0002.0007-1 in the network directory “CertData\426A12 Environmental Test”  for this specific test and store in this directory.

9.5	Unit Failure

If any units did not pass the chamber tests, then troubleshoot the failures, make repairs and repeat the test process.

10.0	RECORDS

The chamber results are maintained electronically for a minimum of two years in the ReportGen database.

11.0	DISTRIBUTION

Manufacturing

12.0	ATTACHMENTS

Not applicable to this procedure.

13.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
|  | A | 9 June 2008 | LWH | Created work instruction |
| 1441 | B | 10 Sept 2013 | LWH | Add making of 426A12 Certificate of Environmental Test D0002.0007-1 |