 RECALIBRATE ALL UNITS TO CUSTOMERS MICROPHONE FOR ANY SERVICE WORK

# CERT-870      {Refer to Work Instruction D0001.8047, D0001.8050, D0001.5051 and D0001.8088}

- Overall visual inspection: (water damage, battery connector, other connectors, etc.)

- On arrival READ and STORE setup parameters, make a Note of the Modem mode if it’s (Yes/No),

also check Baud Rate. Use 870-Utility program these will need to be reentered later on.

- Record temperature compensation, setup parameter 225, and enter this number in the service database.

- Check all keyboard functions.

- Replace Lithium battery.

- Update firmware to current version from the network where applicable.

- Disassemble the 870.

- Check Clock frequency for 32,768.33 Hz adjust if need.

- Inspect board for insects clean if needed.

- Replace input connector if not gold pins.  And /or inspect the input pins and solder connections to the PC-board.

- Reassemble the 870 in preparation to run on the test cart.

- Run the automated test (SLM test).

- Determine Pass/Fail for A&C Weighted frequency results and Log linearity for serial numbers less then 0350.  Serial numbers greater then 0350 need to pass the Crest Factor, Burst and Detector tests.

  - Print out the Data by selecting Mark Cust tab.

    - If the 870 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Run in environmental test chamber.  If requested by customer.

- Print data from report Gen.

- Restore ADDRESS MODE setup parameter 10 back to zero.

- Verify temperature compensation setup parameter 225.  It needs to read the same as recorded in line three, or determined by the environmental test chamber temperature compensation chart.

- When finished testing SEND the STORED setups Modem mode and Baud Rate back into the 870 See line 2

Then hit the store then the Enter key, this will store the setups into the EEPROM.

- Calibrate Model 870 with costumer's pre-amplifier and microphone. If not provided use LD-Certified pre-amplifier and microphone.

- Print Certificate.

- Print Environmental Certificate and record the Cert # on the ENV Certification # doc.

- Print Calibration sticker.

- Print Calibration seal.

- Update LD-data base and technician notes.

- Clean Model 870, before shipping.

## CER-2100         {Refer to Work Instruction D0001.8052 and D0001.5053}

- Remove bird spikes and store in shipping container.

- Remove microphone and Rainhat and deliver it to the microphone laboratory.

- Clean the pre-amplifier

- Remove base tube and inspect for insects around connectors and circuit board.

- Run the automated test (SLM test) select all tests tab.

  - Print out the Data by selecting Mark Cust tab.

    - If the 2100 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Replace O-ring seals, 3900.0013, 3900.0014, 3900.0017 and (3900.0020 * 2).

- Place re-certified microphone on Model 2100.

- Remove Rainhat from Rainhat body and replace O-ring seal 3900.0015.

- Check Rainhat nut that tightens Actuator for proper tightness

- Re-assemble Rainhat.

- Remove any electrostatic charges on Actuator.

- Place Rainhat on microphone and tighten Actuator cable on Rainhat.

- Test Electrostatic Actuator output level for 95.8 dB and 1000 Hz.

- Make sure pre-amplifier Heater works properly, monitor the supply current.  When the heater is turned on with setup parameter 153, the current increases by 50 miliAmperes.

- Run Model 2100 in environmental test chamber.

- Determine PASS/Fail.

- Print Certificate.

- Print Environmental Certificate and record the Cert # on the ENV Certification # doc.

- Calibration sticker.

- Update LD-data base and technician notes.

- Replace desiccants.

- Clean bird spikes and place them back on Model 2100.

- Provide new Windscreen.

## CER-875     {This product is Obsolete work instructions do not apply}

## Overall visual inspection: (water damage, battery connector, other connectors, etc.)

- Disassemble the 875 and Inspect board for insects.

- Replace input connector if not gold pins.  And /or inspect pins solder connections to the PC-board. Two connectors.

- Replace Lithium battery.

- Run on the automated test (Lab test).

  - Print out the Data by selecting Mark Cust tab.

    - If the 875 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Determine Pass/Fail for 1/3 Octave Filters 125 Hz, 1000 Hz, and 8000 Hz and Log linearity.

- Run in environmental test chamber.

- Plot data, determine pass/fail.

- Calibrate 875 with costumer's Model 870, pre-amplifier and microphone.  If not provided use LD-Certified 870, pre-amplifier and microphone.

- Certificate.

- Print Environmental Certificate and record the Cert # on the ENV Certification # doc. Were applicable

- Calibration sticker.

- Calibration seal.

- Update LD-data base and technician notes.

- Clean Model 875, before shipping.

# CER-700    {This product is Obsolete work instructions do not apply}

- Minor REPAIRS CAN BE PERFORMED, THIS IS AN OBSOLETE PRODUCT.

## Overall visual inspection (water damage, battery connector, other connectors, etc.)

- Acoustic sweep on mic/preamp Use the Cal291 for Knowles and EV style 3/8” microphones.

- Run A & C weight on the automated test (LAB test).

  - Print out the Data by selecting Mark Cust tab.

    - If the 700 fail’s print the AS RECEIVED water mark on the Failed data and return to customer.

    - If it passes place the data in with the Certificate.

- Print Certificate

- Cal sticker

- Cal seal

- Calibrate to microphone

- Update database and tech notes

# CER-710   {This product is Obsolete work instructions do not apply}

- Minor REPAIRS CAN BE PERFORMED, THIS IS AN OBSOLETE PRODUCT.

## Overall visual inspection (water damage, battery connector, other connectors, etc.)

- Run A weight filter only on the automated test (LAB test).

  - Test the Black EV style mics and MPR006’s on the automated test (Dosimeter mic preamp test).

  - Use comp #1 for the black EV style and units with Grey EV mics use comp #3 and cannot be certified.

- Print out the Data by selecting Mark Cust tab.

  - If the 710 fail’s print the AS RECEIVED water mark on the Failed data and return to customer.

  - If it passes place the data in with the Certificate.

- Print Certificate

- Cal sticker

- Cal seal

- Calibrate to microphone

- Update database and tech notes

# CER-705   {This product is Obsolete work instructions do not apply}

# Overall visual inspection (water damage, battery connector, etc.)

- Check preamp ground (setscrew)

- Clean preamp and Mic threads

- Acoustic sweep on mic/preamp

- Upgrade firmware

- Run on test station (Lab Test)

  - Print out the Data by selecting Mark Cust tab.

    - If the 705 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Print Certificate

- Cal sticker

- Cal seal

- Calibrate to microphone

## Update database and tech notes

# CER-712/720     {Refer to Work Instruction D0001.8123, D0001.8160 and D0001.8211}

- Overall visual inspection (water damage, battery connector, other connectors, etc.)

- On arrival Read and Store Setups Parameters using the (LD SLM Utility).

- Check preamp ground (setscrew) were applicable.

- Clean preamp and Mic threads were applicable.

- Check all keyboard functions.

- Acoustic sweep on mic/preamp were applicable.

- Update firmware to current version from the network where applicable.

- Run on automated test (Lab Test)

  - Print out the Data by selecting Mark Cust tab.

    - If the 712/720 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- When finished testing SEND the STORED setups back into the instrument.

- Press (Shift + STR, Enter), This will store the setups back into the EEPROM.

- Change the Cal level to 114 dB if tested as MPR005 option.

- Calibrate to the Microphone.

- Print Certificate

- Cal sticker

- Cal seal

- Calibrate to microphone verify offset is correct.

- Update database and tech notes

# CERT-812/820     {Refer to Work Instruction D0001.8123, D0001.8160 and D0001.8211}

- Overall visual inspection (water damage, battery connector, other connectors, etc.)

- On arrival Read and Store Setups Parameters using the (LD SLM Utility).

- Update firmware to current version from the network where applicable.

- Check all keyboard functions.

- Run on automated test (Lab Test)

  - Print out the Data by selecting Mark Cust tab.

    - If the 812/820 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- When finished testing SEND the STORED setups back into the instrument.

- Press (Shift + STR Enter), This will store the setups back into the EEPROM.

- Print Certificate

- Cal sticker

- Cal seal

- Calibrate to microphone verify offset is correct.

- Update database and tech notes

# CER-DSPXX     {Refer to Work Instruction D0001.8145}

- Overall visual inspection (water damage, battery connector, other connectors, etc.)

- We cannot certify any DSPXX with a 3/8” Mic, if one is received in return it AS IS or quote for new unit.

- NO REPAIRS CAN BE PERFORMED on DSPXX with a 3/8” mic.

- DPSXX with a PRM80 or in some cases a PRM828 continue.

- Check all keyboard functions.

- Run on the automated test (Lab Test)

  - Print out the Data by selecting Mark Cust tab.

    - If the DSP fail’s print the AS RECEIVED water mark on the Failed data and return to customer.

    - If it passes place the data in with the Certificate.

- Print Certificate

- Cal sticker

- Cal seal

- Calibrate to microphone

- Update database and tech notes

CER-HVM100    {Refer to Work Instruction D0001.8096, D0001.8097, D0001.8098 and D0001.8099}

- Overall visual inspection (water damage, battery connector, other connectors, etc.)

- If display flickers at start up look for ECO requirements.

- Upgrade boot code to the latest version if applicable.

- Upgrade firmware using (L1, L2, L3, L4) options listed per unit in the data base where applicable.

- Verify the HVM100 memory with a M14 command in hyper terminal, Look for a (FF),

this is an unallocated memory location, if they differ consult Engineering.

- Check all keyboard functions.

- Run on the automated test (HVM Test) select (LogLin, Crest, Burst, Detector, Gain and Fb).

  - Print out the Data by selecting Mark Cust tab.

    - If the HVM100 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Print Certificate

- Cal sticker

- Cal seal

- Update database and tech notes

## CER-SPARK FAMILY   {Refer to Work Instruction under the Spark Family folder on the network}

- Overall visual inspection (water damage, battery connector, other connectors, etc.).

- Check display (706, 706RC, 704) or LED’s (703, 703+, 705 and 705+).

- Check for upgrades (firmware, nylon washers, EMI washers, back label, battery door, etc.).

- Check all keyboard functions (706, 706RC, 704).

- Run the automated test (SLM test) by selecting all tests note: do not select the Pot adjustment box.

- Print out the Data, select Mark Cust only.

- If the SPARK fail’s print the AS RECEIVED water mark on the Failed data.

- If it passes place the data in with the Certificate.

- On older style preamps with electret microphone EV/MIC001/2

  - Run the automated test (Dosimeter mic preamp test).

    - Print data

    - Note: if failed (the MIC Can’t be repaired It will need to be replaced with the New MPR00X style MIC).

- On new style preamps (MPR001/MPR002 with Knowles microphone BL-7046):

  - Run the automated test (Dosimeter mic preamp test).

    - Check the board for electrical tape. (This MIC can be repaired).

    - Print data

- Print Certificate for both the Spark and mic.

- Cal sticker

- Cal seal

- Calibrate to microphone verify offset is correct.

- Update database and tech notes.

## CER-PREAMP     {Refer to Work Instruction under Mics & PreAmps located on the network}

- Overall visual inspection (Pogo, cable connector, tube, etc.).

- Run the automated test (Lab test).

  - Print out the Data by selecting Mark Cust tab.

    - If the PREAMP fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Print Certificate

- Cal sticker

- Calibrate to microphone (check 200v bias voltage) verify offset is correct.

- Update database and tech notes

## CER-2200C   {Refer to Work Instruction D0001.8141}

- Overall visual inspection (water damage, battery connector, other connectors, etc.).

- Check for proper bias voltages (0v, 28v, 100v, 200v).

- Check switch.

- Check ATTN potentiometer.

- Run semi automated test (Lab test).

  - Print out the Data by selecting Mark Cust tab.

    - If the 2200 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Print Certificate

- Cal sticker

- Cal seal

- Calibrate to microphone and verify offset is correct.

- Update database and tech notes

CER-LXT   {Refer to Work Instruction D0001.8305, D0001.8306, D0001.8307, D0001.8308.

- Overall visual inspection (water damage, battery connector, other connectors, etc.).

- Check display

- Upgrade firmware to the most currant version where applicable.

- Check for upgrades (ECO changes or mods etc.).

- Check all keypad functions

- Run the automated test (LxTTEST).

- Select the appropriate tests based on the installed options.

  - Print out the Data using the (report gen) program.

    - (Fails) print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Print Certification

- Cal sticker

- Calibrate to microphone and verify offset is correct.

- Update database and tech notes

CER-831   {Refer to Work Instruction D0001.8309, 8310, 8316, 8317, 8321, 8322, 8334, 8338, 8339

- Overall visual inspection (water damage, battery connector, other connectors, etc.).

- Check display

- Upgrade firmware to the most currant version where applicable.

- Check for upgrades (ECO changes or mods etc.).

- Check all keypad functions

- Run the automated test (LxTTEST).

- Select the appropriate tests based on the installed options.

  - Print out the Data using the (report gen) program.

    - (Fails) print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Print Certification

- Cal sticker

- Calibrate to microphone and verify offset is correct.

- Update database and tech notes

CER-426A12   {Refer to Work Instruction D0001.8326, D0001.8327, D0001.8336

- Remove bird spikes and store in shipping container.

- Remove microphone and Rainhat and deliver it to the microphone laboratory.

- Clean the pre-amplifier

- Remove outer tube and inspect for insects around connectors and circuit board.

- Run the automated test (LAB test).

  - Print out the Data.

    - If the 426A12 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Replace all O-ring.

- Place re-certified microphone on Model 426A12.

- Remove Rainhat from Rainhat body and replace O-rings.

- Check Rainhat nut that tightens Actuator for proper tightness

- Re-assemble Rainhat.

- Remove any electrostatic charges on Actuator (RainHat).

- Place Rainhat on microphone and tighten Actuator cable on Rainhat.

- Test Electrostatic Actuator output level for approximately 96.0 dB.

- Run Model 426A12 in environmental test chamber.

- Determine PASS/Fail, Print out the Data using the (report gen) program.

- Print Certificates.

- Print Environmental Certificate and record the Cert # on the ENV Certification # doc.

- Calibration sticker (Just under the window).

- Replace desiccant and desiccant bag.

- Clean bird spikes and place them back on Model 426A12.

- Provide new Windscreen.

- Update database and tech notes

# CERT-824/814     {Refer to Work Instruction on the network under the 814_824 folder}

- Overall visual inspection (water damage, battery connector, other connectors, etc.)

- On arrival create EEPROM backup ID.

- Update firmware to current version from the network where applicable.

- Verify all ECO’s have been done, complete ones that are not complete.

- Check all keyboard functions.

- Run on automated test (Lab Test)

  - Print out the Data by selecting Mark Cust tab.

    - If the 824/814 fail’s print the AS RECEIVED water mark on the Failed data.

    - If it passes place the data in with the Certificate.

- Run in the environmental Chamber if requested.

- When finished restore Back up ID, press (Check+ REST and ON KEY), this will restore the setups back from the EEPROM to the 824/814.

- Print Certificate

- Skip if not requested on line 7: Print Environmental certificate and record the cert# on the ENV Certification # doc were applicable if requested.

- Cal sticker

- Cal seal

- Calibrate to microphone verify the offset is correct.

- Update database and tech notes

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 1157B | A | 01/25/09 | EO | Initial Release |
| 1308 | B | 09/07/10 | EO | Update for HVM100 added 824/814 and added process changes for all others |