## Thermotron
| Certificate of Calibration | [5 empty] | Certificate No: yyyy-mmdd-x
x next number for that day
xxxxxxx Temperature Humidity Chamber
Model No: xxxxx | [6 empty] | Proceedure: D0001.8343
Serial No: xxxxxx
Test Date: xxxxxxxxx |  |  |  |  | data from test
Due Date: xxxxxxxxx |  |  |  |  | check for latest dates
xxxx specific data entry
Equipment
Sensor: RH Young Temperature Humidity Probe |  |  |  | Model No: 41382
Cert No: 4028079 |  |  |  | Serial N0: 51841140/132228
Cal date: 08/15/08 |  | Recal Date: 08/15/09
Digital Volt Meter: Hewlett Packard |  |  |  | Model No: 34401A
Cert No: 296038 |  |  |  | Serial N0: 3146A73045
Cal date: 01/15/07 |  | Recal Date: 01/15/09
Temperature: 24 °C |  | Humidity: 31% RH
Thermotron |  | Temperature |  |  | Setting | error | Humidity |  | Setting | error
date time | Test | Volts | °F | °C | °C | °C | Volts | % RH | % RH | % RH
2008-08-21 16:01 | REF1 | 0.59567 | 69.1333 | 20.6296 | 20 | 0.62961 | 0.50539 | 50.5391 | 50 | 0.53914
2008-08-21 18:31 | TEMP | 0.14265 | -21.4696 | -29.7054 | -30 | 0.29464
2008-08-21 19:11 | TEMP | 0.23205 | -3.5911 | -19.7728 | -20 | 0.22719
2008-08-21 19:51 | TEMP | 0.32185 | 14.3699 | -9.7945 | -10 | 0.20552
2008-08-21 20:33 | TEMP | 0.41269 | 32.538 | 0.29889 | 0 | 0.29889
2008-08-21 21:13 | TEMP | 0.50259 | 50.5186 | 10.2881 | 10 | 0.2881
2008-08-21 21:53 | TEMP | 0.59136 | 68.2724 | 20.1513 | 20 | 0.15131
2008-08-21 22:33 | TEMP | 0.68227 | 86.4531 | 30.2517 | 30 | 0.2517 | 0.44166 | 44.1664 | 45 | -0.83362
2008-08-21 23:13 | TEMP | 0.77158 | 104.3164 | 40.1758 | 40 | 0.17579 | 0.43324 | 43.3237 | 45 | -1.6763
2008-08-21 23:55 | TEMP | 0.8622 | 122.4411 | 50.245 | 50 | 0.24504 | 0.45394 | 45.3936 | 45 | 0.3936
2008-08-22 00:35 | TEMP | 0.95164 | 140.3276 | 60.182 | 60 | 0.18199 | 0.46377 | 46.3769 | 45 | 1.3769
2008-08-22 01:21 | HUM | 0.77355 | 104.7106 | 40.3948 | 40 | 0.39478 | 0.45373 | 45.3726 | 45 | 0.37255
2008-08-22 01:51 | HUM | 0.77365 | 104.7306 | 40.4059 | 40 | 0.40591 | 0.55379 | 55.3794 | 55 | 0.37938
2008-08-22 02:21 | HUM | 0.77292 | 104.5841 | 40.3245 | 40 | 0.32449 | 0.66022 | 66.0221 | 65 | 1.0221
2008-08-22 02:51 | HUM | 0.77291 | 104.5812 | 40.3229 | 40 | 0.32289 | 0.76429 | 76.4291 | 75 | 1.4291
2008-08-22 03:21 | HUM | 0.77287 | 104.5735 | 40.3186 | 40 | 0.31858 | 0.88322 | 88.3219 | 85 | 3.3219
2008-08-22 03:51 | HUM | 0.77347 | 104.6931 | 40.3851 | 40 | 0.38506 | 0.94838 | 94.8382 | 95 | -0.16178
2008-08-22 04:37 | END | 0.81673 | 113.3455 | 45.192 | 45 | 0.19196 | 0.94866 | 94.8662 | 95 | -0.13375
2008-08-22 05:07 | END | 0.81742 | 113.4848 | 45.2693 | 45 | 0.26932 | 0.94898 | 94.8979 | 95 | -0.10212
2008-08-22 05:39 | END | 0.81747 | 113.4945 | 45.2747 | 45 | 0.27473 | 0.94815 | 94.8154 | 95 | -0.18455
2008-08-22 06:09 | END | 0.81752 | 113.5033 | 45.2796 | 45 | 0.27963 | 0.94812 | 94.812 | 95 | -0.18802
2008-08-22 08:03 | REF2 | 0.59486 | 68.9726 | 20.5403 | 20 | 0.54033 | 0.49841 | 49.8411 | 50 | -0.15888
PCB Piezotronics
Larson Davis a Division of PCB Piezotronics
1681 W 820 N
Leroy Harbaugh |  |  | Provo, Utah 84601
Toll Free: 888.258.3222
Telephone: 716.926.8243
Fax: 716.926.8215
Page 1 of 1 |  |  | ISO 9001-2000 Certified
D0001.8343-1(A) | TEMPERATURE HUMIDITY CHAMBER CERT FORM
Example | Certificate of Calibration | [5 empty] | Certificate No: 2008-0822-1
Thermotron Temperature Humidity Chamber
Model No: SE-1000L | [6 empty] | Proceedure: D0001.8343
Serial No: 36541
Test Date: 22 Aug 2008
Due Date: 22 Aug 2009
Equipment
Sensor: RH Young Temperature Humidity Probe |  |  |  | Model No: 41382
Cert No: 4028079 |  |  |  | Serial N0: 51841140/132228
Cal date: 08/15/08 |  | Recall Date: 08/15/09
Digital Volt Meter: Hewlett Packard |  |  |  | Model No: 34401A
Cert No: 296038 |  |  |  | Serial N0: 3146A73045
Cal date: 01/15/07 |  | Recall Date: 01/15/09
Temperature: 24 °C |  | Humidity: 31% RH
Thermotron |  | Temperature |  |  | Setting | error | Humidity |  | Setting | error
date time | Test | Volts | °F | °C | °C | °C | Volts | % RH | % RH | % RH
2008-08-21 16:01 | REF1 | 0.59567 | 69.1333 | 20.6296 | 20 | 0.62961 | 0.50539 | 50.5391 | 50 | 0.53914
2008-08-21 18:31 | TEMP | 0.14265 | -21.4696 | -29.7054 | -30 | 0.29464
2008-08-21 19:11 | TEMP | 0.23205 | -3.5911 | -19.7728 | -20 | 0.22719
2008-08-21 19:51 | TEMP | 0.32185 | 14.3699 | -9.7945 | -10 | 0.20552
2008-08-21 20:33 | TEMP | 0.41269 | 32.538 | 0.29889 | 0 | 0.29889
2008-08-21 21:13 | TEMP | 0.50259 | 50.5186 | 10.2881 | 10 | 0.2881
2008-08-21 21:53 | TEMP | 0.59136 | 68.2724 | 20.1513 | 20 | 0.15131
2008-08-21 22:33 | TEMP | 0.68227 | 86.4531 | 30.2517 | 30 | 0.2517 | 0.44166 | 44.1664 | 45 | -0.83362
2008-08-21 23:13 | TEMP | 0.77158 | 104.3164 | 40.1758 | 40 | 0.17579 | 0.43324 | 43.3237 | 45 | -1.6763
2008-08-21 23:55 | TEMP | 0.8622 | 122.4411 | 50.245 | 50 | 0.24504 | 0.45394 | 45.3936 | 45 | 0.3936
2008-08-22 00:35 | TEMP | 0.95164 | 140.3276 | 60.182 | 60 | 0.18199 | 0.46377 | 46.3769 | 45 | 1.3769
2008-08-22 01:21 | HUM | 0.77355 | 104.7106 | 40.3948 | 40 | 0.39478 | 0.45373 | 45.3726 | 45 | 0.37255
2008-08-22 01:51 | HUM | 0.77365 | 104.7306 | 40.4059 | 40 | 0.40591 | 0.55379 | 55.3794 | 55 | 0.37938
2008-08-22 02:21 | HUM | 0.77292 | 104.5841 | 40.3245 | 40 | 0.32449 | 0.66022 | 66.0221 | 65 | 1.0221
2008-08-22 02:51 | HUM | 0.77291 | 104.5812 | 40.3229 | 40 | 0.32289 | 0.76429 | 76.4291 | 75 | 1.4291
2008-08-22 03:21 | HUM | 0.77287 | 104.5735 | 40.3186 | 40 | 0.31858 | 0.88322 | 88.3219 | 85 | 3.3219
2008-08-22 03:51 | HUM | 0.77347 | 104.6931 | 40.3851 | 40 | 0.38506 | 0.94838 | 94.8382 | 95 | -0.16178
2008-08-22 04:37 | END | 0.81673 | 113.3455 | 45.192 | 45 | 0.19196 | 0.94866 | 94.8662 | 95 | -0.13375
2008-08-22 05:07 | END | 0.81742 | 113.4848 | 45.2693 | 45 | 0.26932 | 0.94898 | 94.8979 | 95 | -0.10212
2008-08-22 05:39 | END | 0.81747 | 113.4945 | 45.2747 | 45 | 0.27473 | 0.94815 | 94.8154 | 95 | -0.18455
2008-08-22 06:09 | END | 0.81752 | 113.5033 | 45.2796 | 45 | 0.27963 | 0.94812 | 94.812 | 95 | -0.18802
2008-08-22 08:03 | REF2 | 0.59486 | 68.9726 | 20.5403 | 20 | 0.54033 | 0.49841 | 49.8411 | 50 | -0.15888
PCB Piezotronics
Larson Davis a Division of PCB Piezotronics
1681 W 820 N
Leroy Harbaugh |  |  | Provo, Utah 84601
Toll Free: 888.258.3222
Telephone: 716.926.8243
Fax: 716.926.8215
Page 1 of 1 |  |  | ISO 9001-2000 Certified
D0001.8343-1 (A) | TEMPERATURE HUMIDITY CHAMBER CERT FORM