730 Final Assembly and Test

1.0		PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to verify final assembly and test of 730 products.  This process also serves as a verification that the product continues to meet the specifications and criteria established by Engineering.  Instructions apply to both the 730 and 730IS models unless specifically noted.  References to “meter” apply to all products.

Instructions for Intrinsic Safe Model 730IS will be highlighted with this gold shaded background.

2.0		AFFECTED DEPARTMENTS

Manufacturing

3.0		REFERENCE DOCUMENTS

- D0001.8375	METCAL Procedure Instructions

- D0001.8432	Model 730 Final Test, METCAL

- D0001.8439	Model 730 Board Test Instructions

- D0001.8447	Model 730 Mic Test, METCAL

- D0001.1126-1	Quality Records Matrix

- A730.01-IS 	Main Board Schematics and Assembly Drawing

- A730.02 	Lithium Battery Interface Schematic and Assembly Drawing

- S730.01	Case assembly drawing

- S730.02	Lithium battery assembly drawing

- D0001.8441	Model 730IS Final Checklist

- A730.03-IS	NiMH Battery Interface Schematic and Assembly Drawing

- A730.03-IS 	NiMH Battery Interface Schematics and Assembly Drawing

- A730.12-IS	IS Main Board Schematic and Assembly Drawing

- S730.03-IS	Intrinsic Safe NiMH battery assembly drawing

4.0		RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Use proper ESD protective equipment.

- Verify compliance of the product under test to specification.

- Troubleshoot and correct product as required.

- Communicate concerns to Supervisor or Quality Assurance.

- Request management review of product concerns.

5.0		SAFETY PRECAUTIONS

- Safety Glasses are required when soldering, lead clipping, or testing power supplies.

- Follow general electrical precautions for working with energized, low voltage circuits.

6.0		EQUIPMENT & MATERIALS

- ADP106 	ADAPTOR INPUT BNC TO 1/4" MIC 12 pF

- CAL200 	Calibrator

- ADP109	1/4" to 1/2" CALIBRATOR ADAPTER

- TMS9917C-LD 	Microphone Comparison Coupler

- DVX016	BLUETOOTH USB DONGLE

- CBL218	USB cable, A to micro-B (LD PN 0621.0111)

- S730.06	1 to 5 BNC 730 TESTING CABLE

- 831	831 Sound Level Meter

- PRM831	Microphone preamplifier

- Test station 	With signal generator (SRS DS360) and METCAL test software

- Test fixture	730 Multi-unit wireless charging test fixture

- Miscellaneous	Cables, adapters and other miscellaneous test station items

7.0		INSPECTION

With the meter fully assembled as documented in the S730.01 drawing,

- Inspect all meters for functional and cosmetic nonconformities

- Verify all Intrinsically Safe critical parts (A730.12-IS and A730.03-IS), this must be noted on the D0001.8441 form.

- INSTRUCTIONS

- Perform the MET/CAL Final Test

- Connect the ADP106 to the meter(s) microphone input

- Configure Test Station and meter power supply

BNC Cable to

Test Station output

Single unit test configuration:

Power meter with AC–to–USB Adapter and a USB A to Micro-B cable, or with a Qi compatible wireless charger pad.

Connect ADP106 to the Test Station output with a BNC cable.

    - Multiple unit test configuration:

Power meter with Qi compatible wireless multiple charger pad array.

S730.06 Testing Cable to Test Station output

Connect each ADP106 to the Test Station output using the custom S730.06 “730 TESTING CABLE”.

- Ensure that a USB Bluetooth dongle has been installed in the testing PC.

- Power on meter(s).

- Open MET/CAL test software.

- Create asset(s) for the new meter(s).

- Run D0001.8432 730 Final Test from MET/CAL Run Time, follow all prompts.  For multiple units, each test will execute successively and you will only be required to make these selections for the first unit tested.

- Review test results and resolve any issues.

Note: The values in the Range test are calculated based on values taken during Overload, LogLin and Noise Floor tests.  If any of these four tests have to be rerun, all four should be retested to ensure accuracy.

- Remove all connections and power off the meter(s).

- Run the MET/CAL Mic Test (a.k.a. “Acoustic Test” for this meter  (D0001.8375 & D0001.8447)

This test utilizes the CAL200, the TMS9917C-LD Microphone Comparison Coupler, the ADP109 1/4” microphone adapter, the reference microphone provided with the system, the PRM831 and the Model 831 Sound Level Meter.

- Attach the microphone to the meter.

- Ensure that a USB Bluetooth dongle has been installed in the testing PC.

- Connect USB power to meter if meter’s battery power is insufficient to complete test.

- Select D0001.8447 730 Mic Test from MET/CAL Run Time, follow all prompts.

- Review test results and resolve any issues.

- Secure Microphone Attachment

- Secure the microphone to the meter using a very small amount of Blue Thread-Lock on the threads of the meter’s microphone mount (LD PN 0220.0008, LOCTITE THREADLOCKER BLUE 242).

- Wipe away any excess LOCTITE to ensure proper conductivity to microphone.

- Apply Calibration Label

- Install Calibration Label by placing the printed L730.02 calibration label on the back.

- Configure Model Number, Serial Number and Manufacture Date

  - Connect meter to computer via USB

  - Open Spartan 730 Serialize program

    - Select correct COM Port

    - Select the correct Model Type radio button: 730 or 730IS.

    - Enter meter’s Serial Number (failure to enter correct serial number will overwrite number previously programmed into meter)

The Serial Number for the 730IS is entered with the first two digits representing the year manufactured, e.g. 201000 indicates it was manufactured in 2020 and is sequence number 1000.

    - Select correct Mfg. Date

    - Click “Serialize”

  - Navigate to meter’s About view and verify correct model type and serial number information has been programmed into meter

- RECORDS

- Create the D0001.8441 730IS Final Checklist

  - Enter Model as 730IS

  - Enter meter’s serial number

  - Enter microphone’s serial number

  - Enter Tracking Information

- Information entered into the database.

  - Model

  - Serial Number

  - Manufacture Date

  - Certification Date

  - Job Number

  - Firmware Version

  - Software Version

  - Record A730.01-IS and A730.02 board revisions, and mic serial number

- Test data is stored electronically in the 730 database or Metrology database, per the Quality Records Matrix.

- INSPECTION

- Deliver to Logistics

11.0	DISTRIBUTION

This instruction is available to employees electronically via the online Document Control area.

12.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 1876 | A | 13 Mar 2019 | AJR | Initial release |
| 1898 | B | 12 June 2019 | KH & AJR | Moved assembly and board validation info to the S730.01 drawing. Updated instructions for inspection and testing. |
| 1923 | C | 17 Sep 2019 | KH | Added requirement for retesting Range, Overload, LogLin and Noise Floor |
| 1948 | D | 18 Mar 2020 | KH | Added instructions to update manufacture date in meters after certification. |
| 1953 | E | 14 Apr 2020 | AJR | Added 730IS specific items |