# G4 AudCal Kit Configuration Service Work Instruction

#### 1.0	PURPOSE

When servicing AudCal kits (SYS014, SYS015, SYS016 and SYS017), information about the various components included with the kit needs to be entered into a database inside the kit’s 831C.

G4 includes a utility that will write this information into the 831C’s AudCal database.

#### 2.0	AFFECTED DEPARTMENTS

Service

#### 3.0	REFERENCE DOCUMENTS

IG4.01	G4 Software Manual

#### 4.0	RESPONSIBILITIES & AUTHORITY

Authorized users of the software are responsible to…

- Enter the information requested by the utility, which will be written to the AudCal database inside the 831C.

- QA personnel will verify that this information has been entered correctly.

#### 5.0	DEFINITIONS

G4 Command Line – A way to run G4 that enables certain features that are not normally available.

SYS014, SYS015, SYS016, SYS017 – Order numbers for particular configurations of the AudCal Kit.

#### 6.0	SAFETY PRECAUTIONS

No safety precautions indicated for operation of the equipment or performance of this procedure.

#### 7.0	EQUIPMENT & MATERIALS

A desktop PC is required to run the application.

#### 8.0	INSPECTION

There is no inspection associated with this work instruction.

#### 9.0	INSTRUCTION

##### 9.1	Install and Configure G4 Desktop Icon (Only required on new G4 install)

NOTE: The steps in this section will need to be performed by someone with proper administrator privileges.

Install G4 version 3.1.0 or later. To enable the AudCal Kit Configuration utility in G4, add a command line parameter to the icon that will launch G4, as shown in the following steps.

- Right-click the desktop G4 icon and drag it away from the original icon. Release the right-click button. The following menu will pop up:

Click “Copy here” to create a new G4 icon that will be used to run G4 with the configuration script enabled.

- On the new G4 icon, click on the “G4 LD Utility” label and select the “LD”. Replace it with “AudCal” and press Enter. The original icon will run G4 normally.

- Right-click on the new icon and select “Properties”.

| - On the dialog that pops up, edit the Target entry to add -showKitCfg to the end of the entry, after the quotation mark, as shown below. Include a space between the quotation mark and the dash. - Click OK to close the dialog. This icon will now run G4 with the KIT button visible. |  |
| --- | --- |

##### 9.2	Prepare the Frequency Response Files

For microphones certified internally at LD, follow steps 9.2.1. For microphones certified externally (i.e. at PCB), follow steps 9.2.2.

| 9.2.1 Create the frequency response “.csv” file for internally certified microphones using the Mic program as follows: - Open Mic program. - Open folder icon.. - Scroll and select microphone - Scroll and select microphone serial number. If more than one, select data with the latest date.. - Open Sweep results folder. - Click the Make Response File button.. - Select Actuator - Select Ok and close the Mic program - Open shortcut to the Mic icon. File with microphone data should be at the top. Drag this file into: R:\Provo\Logistics\Mic data. |
| --- |
| - To create a “.csv” file for an externally certified microphone, copy the Excel file for this microphone from the folder: R:\Calibration\general\CAL\377SeriesMicrophones to the folder: R:\Provo\Logistics\Mic data Open the Excel file. Save it as a CSV file by selecting the data tab, selecting Save As, then clicking the Browse icon. In the dialog that pops open, click on the drop-down box next to the Save as type label and select the entry CSV (Comma delimited) (*.csv) and click the Save button. Select Okay and Yes for the dialogs that pop up. |  |

    - For an artificial mastoid (AMC493B), copy the frequency response “.csv” file for the artificial mastoid (if included in the kit) from the folder:

R:\Provo\Data\CertData\Mastoid\Automated tests\Export data Mastoids

to the folder:

R:\Provo\Logistics\Mic data

    - Optional – Create CSV File for AEC304 Coupler

- Run the PDF AEC304 Frequency Response Extractor utility to create a frequency response CSV file from the Calibration Certificate PDF file that ships with the coupler.

- Click the Select PDF File button and select the PDF Calibration Certificate file for the AEC304 coupler.

- Click the Select Output Folder button, if necessary, to set the path to the location where the CSV file will be stored. This path will be remembered from previous sessions and may not need to be changed.

- Click the Extract to CSV File button, which will only be enabled after the PDF file has been selected and the output folder specified. The CSV file will be created with a name AEC304_SN_CalDate.csv, such as “AEC304_0139_23-Feb-2018.csv”.

##### 9.4	Ensure Auto-Sync of AudCal Database is turned off in G4

Open your G4 software and follow the instructions below before connecting to a meter.

| Select Tools/Options. |  |
| --- | --- |
| Click on AudCal Options tab. |  |
| Uncheck the checkbox shown. We don’t want the AudCal database to automatically sync when connecting to meter. |  |

##### 9.5	Configure the Kit Using G4

| - Ensure 831C is plugged in to the PC and powered on. - Run G4 with the kit configuration enable by double-clicking the icon modified in step 9.1. |  |  |
| --- | --- | --- |
| - Connect to the 831C from the AudCal kit by selecting it from the meter list. |  |

| - Do not adjust the customer’s time clock on the meter - Click the AudCal mode button if not in AudCal mode. If the AudCal View pane is not full-screen, click in the bar as indicated in the image at right. - After a few seconds, the KIT button will appear as shown at right. - Click the KIT button to show the AudCal Kit Configuration Utility. - Drag the bottom edge of G4 down to show the entire Configuration Utility screen. |  |
| --- | --- |

Note: Leave the fields blank for any item not included in the kit.

- The SLM values will be read from the meter and populated automatically.

- Verify the correct Preamp Model and enter its Serial Number and update the Calibration Due date.

- Verify the correct Coupler Model(s) and enter its Serial Number and update the Calibration Due date.

- Verify correct Microphone(s) is listed. Click Load Data button and select its frequency response file from the local folder. Update the calibration due date.

- Verify correct Artificial Mastoid is listed if included in the system and enter its serial number. Click Load Data button and select its frequency response file from the local folder. Update the calibration due date. Leave the fields blank if a mastoid is not included.

- Verify correct Calibrator Model its serial number are listed and update the calibration due date. Verify that the Freq and Level values are correct.

- If included in the system, click the AEC304 coupler Load Data button and select the CSV file created in section 9.3. The Serial Number will be filled in automatically. Leave the fields blank if the AEC304 coupler is not included.

- Verify that all required fields are filled in correctly.

- Click the Save to Database button.

| - Click the X close button in the meter tab to disconnect the meter. |  |
| --- | --- |

#### 10.0	INSTRUCTIONS FOR REPLACED EQUIPMENT

In the event that equipment, i.e. a microphone, is replaced in the AUDCAL system as part of the equipment service this equipment must be removed and the correct microphone must be re-associated with the coupler.

- While connected to a meter in AUDCAL mode, select My Equipment

- Select the microphones folder

- Select the microphone you wish to remove, and press the red Remove Button. Select Yes

- Click on the coupler and ensure the new microphone is listed in its information.

- Click Save

#### 11.0	USB Documentation

Save the new certifications to the certifications folder on the USB drive. Organize the calibration certifications into their respective years, i.e. all previous years certs go into a folder titled by that year, all current year certs need to go into a folder titled by the current year.

Save the new frequency response files for the microphones and mastoid into the frequency response folder on the USB drive. Note that when a mastoid is calibrated for service, an additional USB drive is typically provided. For service on AUDCAL systems that include a mastoid, write on the job that no USB drive is needed and that the frequency response file should be saved in the ORDERS folder along with the certifications to be copied to the USB drive contained with the system. Organize the files into folders for their respective years, i.e. all previous years files go into a folder titled by that year, all current year files need to go into a folder titled by the current year.

If the customer didn’t return the USB drive with their system, then document D0001.6075 should be followed to send a new drive with the proper documents.

#### 12.0	RECORDS

The only records generated by this work instruction are the database entries on the 831C.

#### 13.0	DISTRIBUTION

This Work Instruction is available online in the work instruction area of Document Control.

#### 14.0	ATTACHMENTS

No attachments applicable to this process.

#### 14.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 1888 | A | 8/15/2019 | DPW | Initial release. |
| 1928 | B | 10/22/2019 | DPW | Added auto-sync and USB with Mastoid instructions |