NIST TRACEABLE 2900 TEST STATION CERTIFICATION

1.0	PURPOSE AND SCOPE

The purpose of this document is to provide the detail and instruction necessary to test and certify the 2900 and 2209/2239 used for our in house test stations.

2.0	AFFECTED DEPARTMENTS

Manufacturing

Quality Assurance

3.0	REFERENCE DOCUMENTS

D0002.0001	Calibration System Controls & Policies

D0001.8397  METTEAM Receiving and Job Creation.doc

4.0	RESPONSIBILITIES & AUTHORITY

The technician has the following responsibilities and authority:

- Verify compliance of the product under test to specification.

- Troubleshoot and correct product as required.

- Communicate concerns to Supervisor or Quality Assurance.

- Request management review of product concerns.

- Use proper ESD protective equipment.

5.0	DEFINITIONS

- When using the 2900, BOLD words represent a hard key function (the keys permanently labeled on the keyboard).

- BOLD [A] words represent a soft key function (the BOLD represents the text on the screen and [A] identifies the hard key associated with the function).

6.0	SAFETY PRECAUTIONS

No additional safety precautions required for this process.

7.0	EQUIPMENT & MATERIALS

- 60db gain box

- BNC to BNC cable(s)

- BNC to mini phone cable

- HP 34401A multimeter (Traceable to NIST)

- RS-232 cable, 9-pin female to 9-pin female null modem

- 12 Volt DC adapter rated to at least 1 Amp

- Two 9 Volt DC adaptors rated to at least 500 mA

- Computer with two com ports, the calibration program, and the old correction file for the system being tested.

- 25-pin male to 25-pin male cable.  This is supplied with the 2209 or 2239.

8.0	INSTRUCTIONS

8.1 Connect all of the equipment

- Connect the HP 34401A multimeter to one of the com ports of the computer.

- Connect the 2900 to the other com port of the computer.

- Plug a 12 Volt adapter into the 2900.

- If the 2900 has a 2209 or 2239 with it, connect the 2209 or 2239 to the 2900’s printer port using a 25-pin male to 25-pin male cable, and plug a 9 Volt adapter into the 2209 or 2239.

- Connect the signal generator out of the 2900 to the appropriate place.

- If there is a 2209/2239 with the system, run the ‘SIGNAL GEN’ out of the 2900 into the input of the 2209/2239, then connect the output of the 2209/2239 to the input of the HP 34401A multimeter.

- If there is not a 2209/2239 with the system, run the ‘SIGNAL GEN’ out of the 2900 into the input of the HP 34401A multimeter.

- Turn on the 2900 (ON, -) and press SYSTEM, I/O[I], 9600[O].

8.2	 Run the test program

- Run the calibration program (C:\CAL\CAL.EXE).

- Go into the equipment setup menu and set the serial ports to the appropriate settings for the equipment connected.

- Enter the serial number of the 2900 in the serial number field.

- Enter the temperature and humidity in the appropriate fields.

- Click on the HP 34401A button and select the serial number of the multimeter you are using.  Also check that the cal date of the multimeter is correct.  NOTE:  If you have to add a multimeter or change the dates for the multimeter, you have to edit the file C:\CAL\LDCALTST.INI.

- Select your name from the Calibrated By pull down list.  If your name is not in the pull down list, you have to edit the file C:\CAL\LDCALTST.INI to add it.

- If the system has been tested before, skip down to section 8.4, frequency sweep.

- If the system is new, continue with section 8.3, build frequency corrections.

8.3  Build frequency corrections file

- If the system has a 2209 with it, make sure the ‘Use 2209’ box is checked.

- If the system has a 2239 with it, make sure the ‘Use 2239’ box is checked.

- If the system does not have a 2209 or 2239 with it, make sure the ‘Use 2209’ and ‘Use 2239’ boxes are not checked.

- Click on ‘Build Frequency Corrections’, and make sure it is the only test selected.

- Click ‘start’.

- Make sure that it starts running.  If it does not, something is set up wrong.

- Test run time will take about an hour.  Test is automated and will not require technician supervision during this part of the test.

- After the test has completed, turn off the ‘Build Frequency Corrections’ button.

8.4	  Frequency sweep

- If the system has a 2209 with it, make sure the ‘Use 2209’ box is checked.

- If the system has a 2239 with it, make sure the ‘Use 2239’ box is checked.

- If the system does not have a 2209 or 2239 with it, make sure the ‘Use 2209’ and ‘Use 2239’ boxes are not checked.

- Click on ‘Frequency Sweep’, and make sure it is the only test selected.

- Click ‘start’.

- Make sure that it starts running.  If it does not, something is set up wrong.

- When asked, enter the serial number of the 2209/2239, or type ‘NONE’ if the system does not have a 2209/2239.

- After the test is over click on ‘VIEW/PRINT RESULTS’ and see if the test passed.

- If the test failed do the following.

- Click on ‘CLOSE’ to exit the results window.

- Turn off the ‘Frequency Sweep’ button.

- Go to step 8.3 ‘Build frequency Corrections’ to create a new correction file then repeat this test.

- If the test passed print the results.

8.5  Amplitude Sweep

NOTE:  This test is only preformed when the test station that is being certified has a 2209 or 2239 with it.  The mic test systems, calibrator test system, MetCal, etc. do not get this test done.

- Plug a 9 Volt adaptor into the 60dB gain box.

- Connect the output of the 2209 or 2239 to the input of the 60dB gain box.

- Connect the output of the 60dB gain box to the input of the Multimeter.

- Make sure that the ‘Use 2209’ or ‘Use 2239’ button is ON (it must be on).

- Click on the ‘Amplitude Sweep’ button, and make sure that it is the only test selected.

- Click ‘start’.

- Follow the directions on the screen.

- After the test is done, click on ‘VIEW/PRINT RESULTS’ and see if the test passed.

- If the test passed print the test results.

8.6  Finish

- Run the 2900 on the computer test station using the Larson Davis RTA test, and print the test results.

- Create a certificate of compliance for the system with a one-year interval.

- If the system has a 2209, create a certificate using 2900 / 2209 as the model.

- If the system has a 2239, create a certificate using 2900 / 2239 as the model.

- If the system does not have a 2209 or 2239, create a certificate using 2900 / SIG GEN as the model.

- Create an in house calibration sticker for the 2900 and place it on the front of the 2900 next to the left edge of the keyboard, and make sure there is a calibration seal on the top, center screw on the back of the 2900.

- If there is a 2209 or 2239, create an in house calibration sticker for the 2209/2239 and place it on the top of the 2209/2239, and make sure there is a calibration seal on the side of the case.

- Update the database for the 2900 and 2209/2239 with the new calibration dates.

- If the ‘Create Frequency Corrections’ test was run, copy the new correction file to the computer that will use that 2900.

- The name of the file is LDxCnnnn.BIN where nnnn is the serial number of the 2900.  x is either ‘A’ for a 2209 system, ‘B’ for a 2239 system, or ‘F’ for a system without a 2209 or 2239.

- These files are located in C:\CAL\

- It goes in the same directory as the database for the SLM test, and in the directory that the program is in for the microphone test program and RTA test.

- If the 2900 is used on a MetCal station, you need to update the MetCal software with the new calibration and due dates.  Follow D0001.8397 METTEAM Receiving and Job Creation.doc

- Give all of the paperwork (certificate, signal generator results, 2900 test results) to Quality Assurance for their records.  If the 2900 that was tested appears on certificates generated out of the old database, let QC know that they need to update the old database.

9.0	RECORDS

Test results are maintained electronically.  Paper copies of the tests are maintained by the QA Manager per the requirements of the "Calibration System Controls & Policies" procedure (D0002.0001).

10.0	DISTRIBUTION

Manufacturing

11.0	ATTACHMENTS

None

12.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 307 | A | 6/15/01 | JGG | Initial Release |
| 1167 | B | 1/27/09 | JGG | Removed GPIB instructions for the Multimeter. Added instructions for systems with a 2239 |
| 1950 | C | 24 Mar 2020 | JGG | Removed cal sticker on back of 2900 and added step to update MetCal. |