HVM100 PRE-SHIPMENT INSPECTION

1.0		PURPOSE AND SCOPE

This document provides the instructions necessary to perform a pre-shipment inspection of the HVM100.

2.0		RESPONSIBILITY

Quality Assurance, or designated person, has the responsibility to follow this instruction when inspecting the products listed in section 1.0.

3.0		MATERIALS AND EQUIPMENT

D0001.2051 HVM100 – HAVPRO Options & Features

Two “AA” batteries

4.0		INSTRUCTIONS

Follow the instructions in this procedure concurrent with the D0001.2045 QA Pre-Shipment Inspection Checklist – All Instruments Instruction.

4.1		Inspection

- Print a new copy of the order verification for new orders.

- Verify the serial number and model of each product against the order and calibration certificates.

- Verify the certification due sticker is on the back of the instrument and the calibration void sticker is on the side.

- Verify the certification date on the instrument matches the calibration certificate.

- Verify the calibration due date is the interval the customer requested on order.  If there is no request, it is left blank.

4.2		New Product Cosmetic Inspection

    - Inspect for gouges, scratches, and dents on the case.

    - Verify the LCD is free of discolorations, scratches, or other blemishes.

    - Verify the mechanical fit of the two halves of the case.  Look for gaps and misalignments.

    - Verify the presence of screws, feet, and Larson Davis jewel.

4.3		Service Product Cosmetic Inspection

1.  Verify the display is clean and free of fingerprints.

4.4		Functional Inspection for New Orders.

- Remove the battery door.

- Install the two batteries per the stickers inside the compartment.

- Verify the battery door fits and clicks in place with the batteries installed.

- Turn on the instrument by pressing the On/Off key.

- Press the “Tools” key.

- Using the up arrow key, arrow up to the “Language” option, and press the center key.

- Using the right arrow, advance through the language choices.  The List below shows the languages that are loaded on each version.  Verify with the technician checklist and order which version the customer ordered.

| Version | Options |
| --- | --- |
| L1 | English, Espanol, Deutsch, Italino |
| L2 | English, Deutsch, Portuguese, Cesky |
| L3 | English, Francais, Italiano, Deutsch |
| L4 | English, Nederlands, Deutsch, Francais |

- When “English” is again displayed, press the center key to select the option.

- Press the “Setup” key.

- Press the center key and use the right arrow key to advance through the setup options and verify it with the technician checklist and order.

| Option | What is included |
| --- | --- |
| No option | Vibration |
| HA | Vibration, Hand-arm |
| WB | Vibration, Whole-body |
| HAWB | Vibration, Hand-arm, Whole-body |
| ALL | Vibration, Hand-arm, Whole-body |

- When “Vibration” is again displayed, press the center key to select the option.

- Press the “On/Off” key.

- Instrument asks if you want to save the new setups, press the left arrow key to respond “NO”.  Instrument will then power off.

- Remove the batteries.

- If the instrument is being shipped on an “Ixxxxxx” order, verify the WEEE symbol is on the back of the instrument.  If no symbol is present, add the WEEE sticker to the back of the instrument.

  - Functional Inspection for Service Orders.

- Remove the battery door.

- Install the two batteries per the stickers inside the compartment.

- Verify the battery door fits and clicks in place with the batteries installed.

- Turn on the instrument by pressing the On/Off key.

- Press the “On/Off” key.

- Instrument asks if you want to save the new setups, press the left arrow key to respond “NO”.  Instrument will then power off.

- Remove the batteries.

- PAPERWORK

Record the inspection on the QA Pre-Shipment Inspection Log, D0002.2044.

For new units, sign a “QA Final Inspection” (D0002.2003) slip and place inside the envelope containing the calibration certificate.

Initial and date the order.  Place the unit on the inspection shelf or notify shipping that the unit is ready.

6.0		REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 1077 | A | 5/09/08 | MRH | Initial release |
| 1789 | B | 2/1/2018 | AC | Update procedure |