HAVPRO PRE-SHIPMENT INSPECTION

1.0		SCOPE AND APPLICABILITY

This document provides the instructions necessary to perform a pre-shipment inspection of the HAVRPO.

2.0		RESPONSIBILITY

Quality Assurance, or designated person, has the responsibility to follow this instruction when inspecting the products listed in section 1.0.

3.0		MATERIALS AND EQUIPMENT

- D0001.2051 HVM100 – HAVPRO Options & Features

- Two “AA” batteries

4.0		INSTRUCTIONS

  - Inspection

The HAVPRO is the HVM100 with a different keypad and color scheme.  Because it is sold and marketed by one of our competitors, the Larson Davis “Calibration Void” sticker is not used.

- Print a new copy of the order verification for new orders

- Verify the serial number and model of each product against the order and calibration certificates.

- Verify the certification due sticker is on the back of the instrument.

- Verify the certification date on the instrument matches the calibration certificate.

- Verify the calibration due date is the interval the customer requested on order.  If there is no request, it is left blank.

- Verify the calibration certificate has been printed on Quest Letterhead.

4.2		New Product Cosmetic Inspection

    - Inspect for gouges, scratches, and dents on the case.

    - Verify the LCD is free of discolorations, scratches, or other blemishes.

    - Verify the mechanical fit of the two halves of the case.  Look for gaps and misalignments.

    - Verify the presence of screws, feet, and red Quest Technologies jewel.

4.3		Service Product Cosmetic Inspection

1.  Verify the display is clean and free of fingerprints.

4.4		FUNCTIONAL INSPECTION

- Remove the battery door.

- Install the two batteries per the stickers inside the compartment.

- Verify the battery door fits and clicks in place with the batteries installed.

- Turn on the instrument by pressing the On/Off key on the keypad.

- Verify Quest comes up on display as the company name on the boot-up screen. (Not Larson Davis)

- Press the “Tools” key.

- Press the arrow up key to the “Language” option, and press the center key.

- Using the right arrow, advance through the language choices—HAVPRO always has L1: English, Espanol, Deutsch, Italiano

- When “English” is again displayed, press the center key to select the option.

- Press the “Setup” key.

- Press the center key and using the left or right arrow keys, advance through the setup options.  The HAVPRO has the –Vibration, Hand Arm, Whole Body options.

- When Vibration is again displayed, press the center key to select the option.

- Press the “On/Off” key.

- Instrument asks if you want to save the new setups, press the left arrow key to respond “NO”.  Instrument will then power off.

- Remove the batteries.

- If the instrument is being shipped on an “Ixxxxxx” order, verify the WEEE symbol is on the back of the HAVPRO.  If no symbol is present, add the WEEE sticker to the back of the instrument.

- PAPERWORK

Record the inspection on the QA Pre-Shipment Inspection Log, D0002.2044.

Initial and date copy of the order.  Place the unit on the inspection shelf or notify shipping that the unit is ready.

6.0		REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 1073 | A | 5/09/08 | MRH | Initial release |
| 1830 | B | 4/27/18 | NJ | Update procedure to print, initial and date copy of order |