NMS044 PRE-SHIPMENT INSPECTION

# 1.0		PURPOSE AND SCOPE

This document provides the instructions necessary to perform a pre-shipment inspection of the NMS044.

# 2.0		RESPONSIBILITY

Quality Assurance, or designated person, has the responsibility to follow this instruction when inspecting the products listed in section 1.0.

# 3.0		MATERIALS AND EQUIPMENT

D0001.2047 831C Pre-shipment Quality Inspection

Windows Computer

For Service – INMS044.05 RV50 NOTICE

# 4.0		INSTRUCTIONS

## 4.1		INSPECTION

- Verify the calibration certificates found on the USB drive under the Calibration folder.

- Verify the serial number.

- Verify the address.

- Verify the cert number.

- Verify the calibration due date is the interval the customer requested on order.  If there is no request, it is left blank.

## 4.2		FUNCTIONAL INSPECTION

- Remove the plate from the case and set the foam block aside.

- Verify the battery is hooked up correctly to the cables, black to negative, red to positive.

- Connect the battery to the power block instead of using the batteries to power the 831C.

- Verify the 831C turns on correctly and follow inspection requirements from D0001.2047 document.

- Verify the RV50 has turned on and the lights are blinking.

- Turn off the system by pressing the On/Off key on the 831C and follow the prompts for shutdown.

- Disconnect battery and place the cable back in the bottom of the case.

- Replace the plate back into the case.

- Replace the foam block between the battery and the plate.

## 4.3 	NEW PRODUCT COSMETIC AND PLACEMENT INSPECTION

    - Inspect for gouges, scratches, and dents on the case as well as the plate.

    - Ensure the screen of the 831C is clean.

    - Verify the DXV015 2 port USB hub is mounted on the top side of the plate above the Genasun charge controller.

    - Verify the O-ring is installed on the knurled 831 retention screw on the top side of the plate.

    - Verify the Anderson Power Pole connectors are plugged into the power block.

    - On the solar panel, verify the Male Pin is Black and the Female Pin is Red.

    - On the CBL228-03 solar cable which comes out of the grommet, Male Pin is Red and Female Pin is Black.  Ensure the plug caps are installed in the end of this cable.

    - Verify the PSA039 Male Pin is Black and the Female Pin is Red.

## 4.4		SERVICE PRODUCT COSMETIC INSPECTION

- Verify the display is clean and free of fingerprints.

- Verify a copy of INMS044.05 RV50 Notice has been included.

## 4.5		SHIPMENT READY VERIFICATION

- Ensure power to the battery has been disconnected from the power block.

- Check the grommet ports for correct threading and damage.  Ensure the plugs are installed on NMS units.

# 5.0	PAPERWORK

Record the inspection on the QA Pre-Shipment Inspection Log, D0002.2044.

For new units, sign a “QA Final Inspection” (D0002.2003) slip and place inside the case.

Initial and date the order.  Place the unit on the inspection shelf or notify shipping that the unit is ready.

# 6.0		REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| ECO 4594 | A | 7/10/2017 | DWA | Initial release |
| 1809 | C | 2/5/2018 | DWA | Revision of physical inspection, modified steps order, added new sections |