QA PRE-SHIPMENT INSPECTION – 820 FAMILY

1.0		PURPOSE AND SCOPE

This document provides the instructions necessary to perform a pre-shipment inspection of the 820 Family including: 712, 720, 812, and 820 instruments.

2.0		RESPONSIBILITY

Quality Assurance, or designated person, has the responsibility to follow this instruction when inspecting the products listed in section 1.0.

3.0		MATERIALS AND EQUIPMENT

- PRM828, if needed.

- ½” LD microphone, if needed.

- Calibrator, if needed.

- 9-Volt battery

4.0		INSTRUCTIONS

Always use the instruments that are being shipped with the product.

4.1		Inspection

- Print a new copy of the Order Verification for new orders.

- Verify the serial number and model of each product against the order and calibration certificates.

- Verify the certification due sticker is on the back of the instrument and the calibration void sticker is on the side.

- Verify the certification date on the instrument matches the calibration certificate.

- Verify the calibration due date is the interval the customer requested on order.  If there is no request, it is left blank.

4.2		New Product Cosmetic Inspection

    - Inspect for gouges, scratches, and dents on the case.

    - Verify the LCD is free of discolorations, scratches, or other blemishes.

    - Verify the mechanical fit of the two halves of the case.  Look for gaps and misalignments.

4.3		Service Product Cosmetic Inspection

1.  Verify the display is clean and free of fingerprints.

4.4		Instructions

Follow the instructions in this procedure concurrent with the D0001.2045 QA Pre-Shipment Inspection Checklist – All Instruments Instruction.

    - Bias Voltage Inspection

If the instrument is not being shipped with a microphone, check the bias voltage to determine which microphone to use to inspect the instrument.

- Apply power to instrument using the 9-volt battery.

- Turn instrument on by pressing “ON” key.

- Press “Setup” key.

- Press “SLM” key.

- Press down arrow 4 times until “Mic Polarization” screen is displayed.

- If the chart says:

| 200 | Use standard microphone |
| --- | --- |
| 0 | Use pre-polarized microphone |

  - Functional Inspection  for New Instruments / Brand New Demo Instruments / Demo Sale Instruments

- Assemble the instrument with preamp, microphone, and calibrator.

- Apply power to instrument using the 9-volt battery.

- Turn instrument on by pressing “ON” key.

- Press “Shift” key, “Reset” key, then “Reset” key to select “Yes”.

- Press “Setup” key.

- Press “SLM” key.

- Detector should be “Slow”.  If not, use right arrow key to select “Slow” and press “R/S” key.

- Press down arrow key.

- Display should be “A”.  If not, use left arrow key to select “A” and press “R/S” key.

- Press “On/Off” key.

- Turn calibrator on.

- Press “Shift” key.

- Press “Cal” key.

- Press down arrow once to start calibration process.  When calibration has been completed, the display will show the words “Done”.

- Verify the offset is between 5 - 22 dB for 820 and 812 and between 23 – 30 dB for 712 and 720

- Record the offset on the technician final checklist, if it hasn’t already been done.

- Record the offset on the QA Pre-shipment Log.

- Press “On/Off” key twice.

- Press “SLM” key.

- Turn calibrator on.

- Press “R/S” key.  The display should read 105.4 dB +/- .5 dB or 114.0 dB + 0.05 dB.  Allow instrument to run for 5 seconds.

- Press “R/S” key.

- Record the cal level on the QA Pre-shipment Log.

- Record reading on the technician final checklist, if it hasn’t already been done.

- Press “Shift” key.

- Press “STR” key.

- Press “R/S” key.

- Press “Shift” key.

- Press “Reset” key.

- Press “Reset” key to select “Yes”.

- Press “Shift” key.

- Press “On/Off” key to turn the instrument off.

- Sign a “QA Final Inspection” (D0002.2003) slip and place inside the envelope containing the calibration certificate.

- Complete the Final Inspection Form (D2140.0018-1), print, sign, and place the form with the 820/812/720/712 certificate.

- If the instrument is being shipped on an “Ixxxxxx” order, verify the WEEE symbol is on the back of the instrument, and the bottom of the preamplifier and microphone box.  If no symbol is found, add the WEEE sticker to the back of the instrument, and the bottom of the microphone and preamplifier box.

4.6  Functional Inspection for Service Instruments / Re-Calibrated Demo Instruments

1.  Assemble the instrument with preamp, microphone, and calibrator.

- Apply power to instrument using the 9-volt battery.

- Turn instrument on by pressing “ON” key.

- Press “Shift” key, “Reset” key, then “Reset” key to select “Yes”.

- Press “Setup” key.

- Press “SLM” key.

- Detector should be “Slow”.  If not, use right arrow key to select “Slow” and press “R/S” key.

- Press down arrow key.

- Display should be “A”.  If not, use left arrow key to select “A” and press “R/S” key.

- Press “On/Off” key.

- Turn calibrator on.

- Press “Shift” key.

- Press “Cal” key.

- Press down arrow once to start calibration process.  When calibration has been completed, the display will show the words “Done”.

- Verify the offset is between 5 - 22 dB for 820 and 812 and between 23 – 30 dB for 712 and 720

- Record the offset on the QA Pre-shipment Log.

- Press “On/Off” key twice.

- Press “SLM” key.

- Turn calibrator on.

- Press “R/S” key.  The display should read 105.4 dB +/- .5 dB or 114.0 dB + 0.05 dB.  Allow instrument to run for 5 seconds.

- Press “R/S” key.

- Record the cal level on the QA Pre-shipment Log.

- Press “Reset” key.

- Press “Reset” key to select “Yes”.

- Press “Shift” key.

- Press “On/Off” key to turn the instrument off.

5.0		PAPERWORK

Record the inspection on the QA Pre-Shipment Inspection Log, D0002.2044.

Initial and date the order.  Place the unit on the inspection shelf or notify shipping that the unit is ready.

6.0		REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
|  | A | 5/26/94 | JEB | Initial Release |
|  | B | 9/23/03 | JEB | Updated to current format |
|  | C | 4/09/07 | JEB | Clarified |
|  | D | 5/09/08 | MRH | Updated inspection procedure |
| 1351 | E | 6/27/11 | MRF | Updated procedure |
| 1378 | F | 4/9/12 | MRF | Updated procedure |
| 1846 | G | 7/10/18 | AC | Updated procedure |