QA FINAL INSPECTION - 800A, 800B

1.0	PURPOSE AND SCOPE

This document provides the instructions necessary to perform a pre-shipment inspection of the 800A and 800B.

2.0	RESPONSIBILITY

Quality Assurance, or designated person, has the responsibility to follow this instruction when inspecting the products listed in section 1.0.

3.0	MATERIALS AND EQUIPMENT

- ¼”, ½”, or 1” Larson Davis microphone, if needed

- PRM826, if needed

- Appropriate adapter for microphone and preamp, if needed

- Calibrator, if needed

- 9-volt adaptor

- Tuning/adjustment tool

4.0	INSTRUCTIONS

Follow the instructions in this procedure concurrent with the D0001.2045 QA Pre-Shipment Inspection Checklist – All Instruments Instruction.

4.1		Inspection

- Print a new copy of the order verification for new orders.

- Verify the serial number and model of each product against the order and calibration certificates.

- Verify the certification due sticker is on the back of the instrument and the calibration void sticker is on the side.

- Verify the certification date on the instrument matches the calibration certificate.

- Verify the calibration due date is the interval the customer requested on order.  If there is no request, it is left blank.

4.2		Service Product Cosmetic Inspection

1.  Verify the display is clean and free of fingerprints.

  - Functional Inspection

- Assemble the microphone, preamp, and calibrator using the appropriate adaptor, if needed.

- Verify the settings on right side of instrument read from top to bottom:  80-140; Lin; Cont; Off; Fast.

- Apply power to the instrument using the 9-volt adaptor.

- Turn calibrator on.

- Press “Power On” key on instrument.

- Unit should stabilize at 114.0 dB +/- .01dB.If it doesn’t, use the adjustment tool to turn the adjustment pot located on the right side of the nose cone until the value is obtained.   Record this reading on the QA inspection log D0002.2044.

- Turn calibrator on.

- Use the adjustment tool to push the button on the center front of the nose cone.  The level should read 134.0 dB +/- 0.1 dB.  If it doesn’t, use the adjustment tool to turn the pot located on the left center of the nose cone until the value is obtained.  The value should hold for 30 seconds.

- Press the “Power On/Off” key.

- Unplug the 9-volt power adaptor from the instrument before removing the microphone and preamp from the instrument.

- If the instrument is being shipped on an “Ixxxxxx” order, verify the WEEE symbol is on the back of the instrument and the bottom of the preamplifier and microphone box.  If no symbol is found, add the WEEE sticker to the back of the instrument and the bottom of the microphone and preamplifier box.

5.0	PAPERWORK

Date and initial the order.   Place the unit on the inspection shelf or notify shipping that the unit is ready.

6.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 345A | B | 10/30/01 | JEB | Rewrite of a very old inspection document for a now obsolete product. Re-write was done to enable inspection of customer instruments prior to shipment. |
| 1075 | C | 5/09/08 | MRH | Updated procedure |
| 1817 | D | 4/25/18 | NJ | Updated procedure |