LXT/831 LCD GLASS INSPECTION

- SCOPE

The purpose of this work instruction is to define the inspection criteria to be used on the LXT/831 LCD glass (p/n M770.0005).

- APPLICABILITY

Quality Assurance

Production

Materials Control

- RESPONSIBILITY

Quality Assurance or a designated representative is responsible for the inspection of LXT/831 LCD glass at incoming inspection.  Inspection will be to the requirements of the fabrication drawing and this work instruction.

- MATERIALS, TOOLS, & EQUIPMENT

Digital calipers or optical comparator

Inspection light with magnification

Glass cleaner and towelette for glass

Reference samples of acceptable and non-acceptable conditions

- INSTRUCTIONS

Inspect under full spectrum fluorescent lights.

Viewing Distance:  15”-24”.  Magnification to be used to confirm suspected defects.

Parts must be examined at a variety of angles to simulate actual use conditions.

Verify dimensions and requirements as detailed on fabrication drawing.

Verify window is clear and free from scratches, chips, blemishes and other defects.

Exception:  Each sheet is allowed a maximum of 1 coating void (appears as a purple dot) of .005” diameter maximum.  Reference samples provided.

Rejects due to silk-screening errors and blemishes on the glass should be returned to vendor.

- RECORDS

Records are maintained for parts that are rejected using the “Disposition of Non-Conformant Materials” (D0001.1034) procedure.

- DISTRIBUTION

This instruction is available to employees in the Quality work instruction folder in the online Document Control area.

8.0	REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
| 892 | A | 9/21/05 | CT | Initial Release |
| 1787 | B | 1/29/18 | AC | Add instruction and correct typographical errors. |