CALIPER CERTIFICATION INSTRUCTIONS

1.0		SCOPE & APPLICABILITY

Internal certification of calipers for in-house use.

2.0		RESPONSIBILITY

Primary responsibility for the implementation of this instruction belongs to the QA Manager and the individual(s) assigned to carry out the process.

3.0		MATERIALS, TOOLS, & EQUIPMENT

Caliper Master

D0002.2053 Internal Caliper Certification form

4.0		INSTRUCTIONS

4.1		Making the Measurements

- Carefully remove the Caliper Master from its protective box and place on a level work surface.

- Verify that the Master is still within its calibration interval.

- Record the Caliper & Master information on the certification form (D0002.2053)

- Using the large jaws (used for outside measurements) measure each of the referenced distances and record the measurement obtained with the test calipers.

- Using the small jaws (used for internal measurements) measure each of the referenced distances and record the measurement obtained with the test calipers.

- Calculate the variation (+/-) of each measurement. (Reference – Measured)

- Record the values in the variations column for each calculation

- Fill in the required information on the Calibration sticker and apply to the back of the calipers or on the carrying case for the instrument.

4.2		Certification Interval

| Caliper Use | Interval |
| --- | --- |
| Quality Assurance Inspections | 3 months |
| Production / Manufacturing | 12 months |
| Engineering | 12 months |

5.0		RECORDS

The completed Caliper Certification form is retained by the QA Manager for at least 3 years after the date of certification.

6.0		DISTRIBUTION

This instruction is available in the on line Document Control area.

7.0		REVISION HISTORY

| DCO # | REV | DATE | INITIALS | CHANGES MADE |
| --- | --- | --- | --- | --- |
|  | A | 6/5/07 | JEB |  |
| 1196 | B | 4/29/09 | MRF | Initial Release |