## TC012
[5 empty] | TC012 | Pressure Transducer Acceleration Sensitivity Worksheet | [7 empty] | Notes: | [5 empty] | Special cases:
[8 empty] | Click for Instructions → | TC012i | [11 empty] | 1807-01 is graded C & D. Models using A & B grade are obsolete.
48537-02 does not get tested for acceleration sensitivity or resonant frequency
Item # | START HERE |  |  | Trim Range |  | mV/g |  |  | Phase |  | Degrees | [9 empty] | 21899-01 calls out a maximum capacitence on the spec drawing. This affected the formula in column AF.
Job #: |  |  |  | to |  | mV/g |  |  | Acceleration Compensation |  | Inches / mV/G
| *All yellow cells must be filled in |  |  | Trim Target |  | mV/g |  |  | Max Trim |  | Inches
| Initial | Calibration | Machining | 2nd | Calibration | Machining | 3rd | Calibration | Machining | Final | Calibration | Specifications (Reference Only) | [6 empty] | Grading
| By: |  | - | By: |  | - | By: |  | - | By: | [5 empty] | 0 | 0 |  | 1
| Date: |  | - | Date: |  | - | Date: |  | - | Date: | [5 empty] | 0 | 0 |  | 2
| Station: |  |  | Station: |  |  | Station: |  |  | Station: |  |  | to | to |  | 0 | 0 |  | 3 |  |  |  | 21899-01
| Amp Cap: |  |  | Amp Cap: |  |  | Amp Cap: |  |  | Amp Cap: | [5 empty] | 0 | 0 |  | 4 | Missing Specs? | TRUE |  |  |  |  | TRUE | FALSE | FALSE |  |  | Acceleration Compensation
Symbol # | Accel Output (mV/g) | Phase Angle (Degrees) | Trim Amount (Inches) | Accel Output (mV/g) | Phase Angle (Degrees) | Trim Amount (Inches) | Accel Output (mV/g) | Phase Angle (Degrees) | Trim Amount (Inches) | Accel Output (mV/g) | Phase Angle (Degrees) | Insulation Resistance (Ω) | Element Capacitence (pF) | Charge Cal (pC/PSI) | Resonant Frequency (kHz) | Linearity (%) | Acceleration Sensitivity (PSI / g) | Element Grade | Data Missing? | Missing Specs? | Insulation Resistance (Ω) | Minumum Capacitence (pF) | Maximum Capacitence (pF) | Minumum Charge Cal (pC/PSI) | Maximum Charge Cal (pC/PSI) | Resonant Frequency (kHz) | Acceleration Sensitivity (PSI / g) | Linearity (%) | Fail? | Element Grade | By: | By: | By: | Output after last trim
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[19 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
TC012 | Rev.J |  | 2020-04-09
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Note on AE14: Anthony J. Agnello: Formula is omitted if Specs are missing. Formula omitted for all item numbere except 21899-01.
Note on AF14: Anthony J. Agnello: Formula is omitted if Specs are missing. Formula omitted for all item numbere except 21899-01.
Note on AG14: Anthony J. Agnello: Formula is omitted if Specs are missing.
Note on AH14: Anthony J. Agnello: Formula is omitted if Specs are missing.

## Specifications
ALL VALUES FOR REFERENCE ONLY, REFER TO SPEC SHEET FOR ACCURACY
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30
Item # | Trim Range |  | Trim Target | Phase | Acceleration Compensation |  |  | Max Trim |  |  |  |  | No-Grade |  |  | A-Grade |  |  | B-Grade |  |  | C-Grade |  |  | D-Grade |  |  | Charge cal |  |  |  | Entered / Updated:
| LSL | USL | [6 empty] | Final Accel | Minimum Insulation Resistance | Minimum Capacitence | Maximum Capacitence | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Minimum Charge cal | Maximum Charge cal | Nominal Capacitence | Missing specs? | By: | Date: | Notes:
| mV/g | mV/g | mV/g | Degrees | Inches | mV/G | Inches / mV/G | Inches | mV | Ω | pF | pF | kHz | PSI / g | % | kHz | PSI / g | % | kHz | PSI / g | % | kHz | PSI / g | % | kHz | PSI / g | % | pC/PSI | pC/PSI
14809-06 | 0 | 0.08 | 0.04 | 180 | 0.005 | 0.05 | 0.1 | 0.014 | 3 | 1000000000000 | 7.5 | 12.5 |  |  |  | 400 | 0.005 | 2 | [11 empty] | 10 | FALSE | Bassano | 2020-04-03
14907-01 | 0.09 | 0.155 | 0.1225 | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-03-31
14907-02 | 0.25 | 0.35 | 0.3 | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-03-31
14907-03 | 0.09 | 0.155 | 0.1225 | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-03-31
14907-08 | 0.25 | 0.35 | 0.3 | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-04-03
14908-01 | 0.09 | 0.155 | 0.1225 | 0 | 0.003 | 0.05 | 0.06 | 0.014 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-04-03
14908-02 | 0.175 | 0.225 | 0.2 | 0 | 0.003 | 0.05 | 0.06 | 0.014 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-04-03
17468-01 | 0.65 | 0.9 | 0.775 | 180 | 0.001 | 0.05 | 0.02 | 0.03 | 12 | 1000000000000 | not tested | not tested |  |  |  | Not Tested | Not tested | not tested | [11 empty] | Not Tested | FALSE | Beal | 2020-04-01
1807-01 | 0.05 | 0.125 | 0.0875 | 180 | 0.005 | 0.05 | 0.1 | 0.014 | 1 | 1000000000000 | 7.5 | 15 |  |  |  | NA | NA | NA | NA | NA | NA | 400 | 0.002 | 0.4 | 400 | 0.002 | 1 | 0.312 | 0.468 | 11.25 | FALSE | Blake | 2020-04-03 | NEED "D" GRADE COLUMN - DONE
1807-02 - OBSOLETE | 0 | 0.1 | 0.05 | 0 | 0.005 | 0.05 | 0.1 | 0.014 | 1 | 1000000000000 | 7.5 | 15 |  |  |  | OBS | OBS | OBS | [11 empty] | 11.25 | FALSE |  |  | No historical data/specs
21899-01 | 0.25 | 0.35 | 0.3 | 180 | 0.003 | 0.05 | 0.06 | 0.01 | 2 | 1000000000000 | 14 | 28 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 21 | FALSE | Beal | 2020-04-01
342-0625-01 | 0 | 0.2 | 0.1 | 0 | 0.003 | 0.05 | 0.06 | 0.014 | 3 | 1000000000000 | 17 | 24 |  |  |  | 200 | 0.01 | 1 | [9 empty] | 0.99 | 1.375 | 20.5 | FALSE | Blake | 2020-03-26
342-0625-02 - OBSOLETE | 0 | 0.2 | 0.1 | 180 | 0.003 | 0.05 | 0.06 | 0.014 | 3 | 1000000000000 | 17 | 24 |  |  |  | 250 | 0.002 | 1 | 200 | 0.002 | 1 | [6 empty] | 0.99 | 1.375 | 20.5 | FALSE | Blake | 2020-03-26
40520-01 | 0 | 0.055 | 0.0375 | 180 | 0.005 | 0.03 | 0.16667 | 0.014 | 1 | 1000000000000 | 9 | 15 |  |  |  | 500 | 0.002 | 0.4 | 500 | 0.002 | 1 | 400 | 0.002 | 2 |  |  |  | 0.3315 | 0.4485 | 12 | FALSE | Kopra | 2020-03-26
40520-02 | 0 | 0.075 | 0.0375 | 180 | 0.005 | 0.03 | 0.16667 | 0.014 | 1 | 1000000000000 | 7.5 | 12.5 |  |  |  | 500 | 0.002 | 0.4 | 500 | 0.002 | 1 | 400 | 0.002 | 2 | [5 empty] | 10 | FALSE | Kopra | 2020-03-26
40726-01 | 0 | 0.075 | 0.0375 | 0 | 0.005 | 0.05 | 0.1 | 0.014 | 1 | 1000000000000 | 7.5 | 12.5 |  |  |  | 500 | 0.002 | 1 | 300 | 0.01 | 2 | [8 empty] | 10 | FALSE | Kopra | 2020-03-31
44986-01 | 0 | 0 | 0 | NA | NA | NA | NA | NA | 0 | 1000000000000 | 7.5 | 12.5 |  |  |  | 0 | 0 | 0 | [11 empty] | 10 | FALSE | Beal | 2020-04-01
46282-01 | 0.15 | 0.25 | 0.2 | 0 | 0.001 | 0.035 | 0.028571 | 0.03 | 8 | 100000000000 | 22.5 | 37.5 |  |  |  | 55 | 0.002 | 1 | [9 empty] | 5.1 | 6.9 | 30 | FALSE | Kopra | 2020-03-31
48537-02 | Not Tested | Not Tested | #VALUE! | Not Tested | NA | NA | #VALUE! | NA | NA | 1000000000000 | 14 | 28 |  |  |  | Not Tested | Not tested | 2 | [9 empty] | 0.88 | 1.32 | 21 | FALSE | Blake | 2020-04-03 | Initial & final accel not tested, res freq not tested
5080-01 | 0.2 | 0.3 | 0.25 | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 25 | 0.002 | 1 | [9 empty] | 0.88 | 1.32 | 20 | FALSE | Blake | 2020-04-02
5080-02 - OBSOLETE | 0.85 | 0.95 | 0.9 | 180 | 0.003 | 0.05 | 0.06 | 0.03 | 2 | 1000000000000 | 15 | 25 |  |  |  | 28 | 0.002 | 2 | [9 empty] | 0.88 | 1.32 | 20 | FALSE | Blake | 2020-04-02
56646-01 | 0 | 0.075 | 0.0375 | 180 | 0.005 | 0.05 | 0.1 | 0.014 | 1 | 1000000000000 | 7.5 | 12.5 |  |  |  | 500 | 0.002 | 0.4 | [11 empty] | 10 | FALSE | Kopra | 2020-03-31
57986-01 | 0.09 | 0.155 | 0.1225 | 180 | 0.003 | 0.05 | 0.06 | 0.014 | 2 | 1000000000000 | 13.5 | 22.5 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 18 | FALSE | Kopra | 2020-04-04 | We have never built these
57986-02 | 0.25 | 0.35 | 0.3 | 180 | 0.003 | 0.05 | 0.06 | 0.014 | 2 | 1000000000000 | 13.5 | 22.5 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 18 | FALSE | Kopra | 2020-04-04 | Used for a fixture - built once in 2015
62903-01 | 0.5 | 0.75 | 0.625 | 180 | 0.001 | 0.05 | 0.02 | 0.02 | 20 | 1000000000000 | 19.5 | 32.5 |  |  |  | Not Tested | Not tested | not tested | [11 empty] | 26 | FALSE | Beal | 2020-04-01
67400-01 | 0.2 | 0.45 | 0.325 | 0 | 0.003 | 0.05 | 0.06 | 0.014 | 3 | 1000000000000 | 17 | 24 |  |  |  | 200 | 0.003 | 1 | [9 empty] | 0.935 | 1.265 | 20.5 | FALSE | Blake | 2020-03-26
67400-02 - PRELIMINARY |  |  | 0 |  | 0.003 | 0.05 | 0.06 | 0.014 | 3 | 1000000000000 | 17 | 24 |  |  |  | 200 | 0.003 | 1 | [11 empty] | 20.5 | FALSE | Blake | 2020-03-26
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
TRUE
Note on E5: Anthony J. Agnello: In Pahse = 0° Out of Phae = 180°
Note on A21: Placebo Element, should have no output.
Note on AC22: Christopher Kopra: Unsure of this . . . Might be for 116B - which the 46282 references in the router
Note on C26: Christopher Kopra: Need to update Router - still in mV

## Engineering Analysis
START HERE |  | START HERE, 0
0
| Acceleration Compensation |  |  | Diaphragm Shift
Symbol # | By: | By: | By: | Output after last trim | Final Output
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0 | *All yellow cells must be filled in
| #N/A | #N/A | #N/A |  | 0 | Initial
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.