## TC012
[5 empty] | TC012 | Pressure Transducer Acceleration Sensitivity Worksheet | [7 empty] | Notes:
[8 empty] | Click for Instructions → | TC012i | [12 empty] | Initial Trim Range |  | mV/g |  | Trim Range |  | mV/g
[22 empty] | to |  | mV/g |  | to |  | mV/g
Item # | START HERE |  |  |  | Min. Output Before Trim |  | mV/g |  |  | Acceleration Compensation |  | Inches / mV/G | [9 empty] | Initial Trim Target |  | mV/g |  | Trim Target |  | mV/g
Job #: |  |  |  |  | Phase |  | Degrees |  |  | Max Trim |  | Inches
| *All yellow cells must be filled in |  |  | *Cells that turn orange are outside of the expected values.
[23 empty] | Special cases: | 1807-01 is graded C & D. Models using A & B grade are obsolete.
| Initial | Calibration | Machining | 2nd | Calibration | Machining | 3rd | Calibration | Machining | Final | Calibration | Specifications (Reference Only) | [8 empty] | Grading |  |  | 48537-02 does not get tested for acceleration sensitivity or resonant frequency
| By: |  | - | By: |  | - | By: |  | - | By: | [5 empty] | 0 | 0 |  |  |  | 1 |  |  | 21899-01 calls out a maximum capacitence on the spec drawing. This affected the formula in column AF.
| Date: |  | - | Date: |  | - | Date: |  | - | Date: | [5 empty] | 0 | 0 |  |  |  | 2
| Station: |  |  | Station: |  |  | Station: |  |  | Station: |  |  | to | to |  | 0 | 0 |  |  |  | 3 |  |  |  | 21899-01
| Amp Cap: |  |  | Amp Cap: |  |  | Amp Cap: |  |  | Amp Cap: | [5 empty] | 0 | 0 |  |  |  | 4 | Missing Specs? | TRUE |  |  |  |  | TRUE | FALSE | FALSE |  |  | Acceleration Compensation
Symbol # | Accel Output (mV/g) | Phase Angle (Degrees) | Trim Amount (Inches) | Accel Output (mV/g) | Phase Angle (Degrees) | Trim Amount (Inches) | Accel Output (mV/g) | Phase Angle (Degrees) | Trim Amount (Inches) | Accel Output (mV/g) | Phase Angle (Degrees) | Insulation Resistance (Ω) | Element Capacitence (pF) | Charge Cal (pC/PSI) | Resonant Frequency (kHz) | Linearity (%) | Acceleration Sensitivity (PSI / g) | Element Grade | Failure Reason code | Failure Reason code | Data Missing? | Missing Specs? | Insulation Resistance (Ω) | Minumum Capacitence (pF) | Maximum Capacitence (pF) | Minumum Charge Cal (pC/PSI) | Maximum Charge Cal (pC/PSI) | Resonant Frequency (kHz) | Linearity (%) | Acceleration Sensitivity (PSI / g) | Fail? | Element Grade | By: | By: | By: | Output after last trim
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
[21 empty] | TRUE | TRUE | [8 empty] | FALSE | 0
TC012 | Rev.K |  | 2020-07-14
| 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Note on AG14: Anthony J. Agnello: Formula is omitted if Specs are missing. Formula omitted for all item numbere except 21899-01.
Note on AH14: Anthony J. Agnello: Formula is omitted if Specs are missing. Formula omitted for all item numbere except 21899-01.
Note on AI14: Anthony J. Agnello: Formula is omitted if Specs are missing.
Note on AJ14: Anthony J. Agnello: Formula is omitted if Specs are missing.

## Specifications
ALL VALUES FOR REFERENCE ONLY, REFER TO SPEC SHEET FOR ACCURACY
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39
Item # | Trim Range |  | Trim Target | Initial Trim Range |  | Initial Trim Target | Min. Output Before Trim | Phase | Acceleration Compensation |  |  | Max Trim |  |  |  |  | No-Grade |  |  | A-Grade |  |  | B-Grade |  |  | C-Grade |  |  | D-Grade |  |  | Charge cal |  |  |  | Entered / Updated:
| LSL | USL |  | LSL | USL | [7 empty] | Final Accel | Minimum Insulation Resistance | Minimum Capacitence | Maximum Capacitence | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Resonant Frequency (Minimum) | Acceleration Sensitivity | Linearity | Minimum Charge cal | Maximum Charge cal | Nominal Capacitence | Missing specs? | By: | Date: | Notes:
| mV/g | mV/g | mV/g | mV/g | mV/g | mV/g | mV/g | Degrees | Inches | mV/G | Inches / mV/G | Inches | mV | Ω | pF | pF | kHz | PSI / g | % | kHz | PSI / g | % | kHz | PSI / g | % | kHz | PSI / g | % | kHz | PSI / g | % | pC/PSI | pC/PSI
14809-06 | 0 | 0.08 | 0.04 |  |  |  | n/a | 180 | 0.005 | 0.05 | 0.1 | 0.014 | 3 | 1000000000000 | 7.5 | 12.5 |  |  |  | 400 | 0.005 | 2 | [11 empty] | 10 | FALSE | Bassano | 2020-04-03
14907-01 | 0 | 0.155 | 0.1225 |  |  |  | 0.2 | 180 | 3 | 0.05 | 60 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-03-31
14907-02 | 0.25 | 0.35 | 0.3 |  |  |  | 0.2 | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-03-31
14907-03 | 0 | 0.155 | 0.1225 |  |  |  | 0.2 | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-03-31
14907-08 | 0.25 | 0.35 | 0.3 |  |  |  | n/a | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-04-03
14908-01 | 0.09 | 0.155 | 0.1225 |  |  |  | n/a | 0 | 0.003 | 0.05 | 0.06 | 0.014 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-04-03
14908-02 | 0.175 | 0.225 | 0.2 |  |  |  | n/a | 0 | 0.003 | 0.05 | 0.06 | 0.014 | 2 | 1000000000000 | 15 | 25 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 20 | FALSE | Bassano | 2020-04-03
17468-01 | 0.65 | 0.9 | 0.775 |  |  |  | n/a | 180 | 0.001 | 0.05 | 0.02 | 0.03 | 12 | 1000000000000 | not tested | not tested |  |  |  | Not Tested | Not tested | not tested | [11 empty] | Not Tested | FALSE | Beal | 2020-04-01
1807-01 | 0.05 | 0.125 | 0.0875 |  |  |  | n/a | 180 | 0.005 | 0.05 | 0.1 | 0.014 | 1 | 1000000000000 | 7.5 | 15 |  |  |  | NA | NA | NA | NA | NA | NA | 400 | 0.002 | 0.4 | 400 | 0.002 | 1 | 0.312 | 0.468 | 11.25 | FALSE | Blake | 2020-04-03 | NEED "D" GRADE COLUMN - DONE
1807-02 - OBSOLETE | 0 | 0.1 | 0.05 |  |  |  | n/a | 0 | 0.005 | 0.05 | 0.1 | 0.014 | 1 | 1000000000000 | 7.5 | 15 |  |  |  | OBS | OBS | OBS | [11 empty] | 11.25 | FALSE |  |  | No historical data/specs
21899-01 | 0.25 | 0.35 | 0.3 |  |  |  | n/a | 180 | 0.003 | 0.05 | 0.06 | 0.01 | 2 | 1000000000000 | 14 | 28 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 21 | FALSE | Beal | 2020-04-01
342-0625-01 | 0 | 0.2 | 0.1 |  |  |  | n/a | 0 | 0.003 | 0.05 | 0.06 | 0.014 | 3 | 1000000000000 | 17 | 24 |  |  |  | 200 | 0.01 | 1 | [9 empty] | 0.99 | 1.375 | 20.5 | FALSE | Blake | 2020-03-26
342-0625-02 - OBSOLETE | 0 | 0.2 | 0.1 |  |  |  | n/a | 180 | 0.003 | 0.05 | 0.06 | 0.014 | 3 | 1000000000000 | 17 | 24 |  |  |  | 250 | 0.002 | 1 | 200 | 0.002 | 1 | [6 empty] | 0.99 | 1.375 | 20.5 | FALSE | Blake | 2020-03-26
40520-01 | 0 | 0.055 | 0.0375 | 0 | 0.04 |  | 0.06 | 180 | 0.005 | 0.03 | 0.16667 | 0.014 | 1 | 1000000000000 | 9 | 15 |  |  |  | 500 | 0.002 | 0.4 | 500 | 0.002 | 1 | 400 | 0.002 | 2 |  |  |  | 0.3315 | 0.4485 | 12 | FALSE | Kopra | 2020-03-26
40520-02 | 0 | 0.075 | 0.0375 |  |  |  | n/a | 180 | 0.005 | 0.03 | 0.16667 | 0.014 | 1 | 1000000000000 | 7.5 | 12.5 |  |  |  | 500 | 0.002 | 0.4 | 500 | 0.002 | 1 | 400 | 0.002 | 2 | [5 empty] | 10 | FALSE | Kopra | 2020-03-26
40726-01 | 0 | 0.055 | 0.0375 | 0 | 0.04 |  | 0.06 | 0 | 0.005 | 0.03 | 0.16667 | 0.014 | 1 | 1000000000000 | 7.5 | 12.5 |  |  |  | 500 | 0.002 | 1 | 300 | 0.01 | 2 | [8 empty] | 10 | FALSE | Kopra | 2020-03-31
44986-01 | 0 | 0 | 0 |  |  |  | n/a | NA | NA | NA | #VALUE! | NA | 0 | 1000000000000 | 7.5 | 12.5 |  |  |  | 0 | 0 | 0 | [11 empty] | 10 | FALSE | Beal | 2020-04-01
46282-01 | 0.15 | 0.25 | 0.2 |  |  |  | 0.25 | 0 | 0.001 | 0.035 | 0.028571 | 0.03 | 8 | 100000000000 | 22.5 | 37.5 |  |  |  | 55 | 0.002 | 1 | [9 empty] | 5.1 | 6.9 | 30 | FALSE | Kopra | 2020-03-31
48537-02 | Not Tested | Not Tested | #VALUE! |  |  |  | n/a | Not Tested | NA | NA | #VALUE! | NA | NA | 1000000000000 | 14 | 28 |  |  |  | Not Tested | Not tested | 2 | [9 empty] | 0.88 | 1.32 | 21 | FALSE | Blake | 2020-04-03 | Initial & final accel not tested, res freq not tested
5080-01 | 0.2 | 0.3 | 0.25 |  |  |  | n/a | 180 | 0.003 | 0.05 | 0.06 | 0.015 | 2 | 1000000000000 | 15 | 25 |  |  |  | 25 | 0.002 | 1 | [9 empty] | 0.88 | 1.32 | 20 | FALSE | Blake | 2020-04-02
5080-02 - OBSOLETE | 0.85 | 0.95 | 0.9 |  |  |  | n/a | 180 | 0.003 | 0.05 | 0.06 | 0.03 | 2 | 1000000000000 | 15 | 25 |  |  |  | 28 | 0.002 | 2 | [9 empty] | 0.88 | 1.32 | 20 | FALSE | Blake | 2020-04-02
56646-01 | 0 | 0.075 | 0.0375 |  |  |  | n/a | 180 | 0.005 | #VALUE! | #VALUE! | 0.014 | 1 | 1000000000000 | 7.5 | 12.5 |  |  |  | 500 | 0.002 | 0.4 | [11 empty] | 10 | FALSE | Kopra | 2020-03-31
57986-01 | 0.09 | 0.155 | 0.1225 |  |  |  | n/a | 180 | 0.003 | 0.05 | 0.06 | 0.014 | 2 | 1000000000000 | 13.5 | 22.5 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 18 | FALSE | Kopra | 2020-04-04 | We have never built these
57986-02 | 0.25 | 0.35 | 0.3 |  |  |  | n/a | 180 | 0.003 | 0.05 | 0.06 | 0.014 | 2 | 1000000000000 | 13.5 | 22.5 |  |  |  | 250 | 0.002 | 1 | [11 empty] | 18 | FALSE | Kopra | 2020-04-04 | Used for a fixture - built once in 2015
62903-01 | 0.5 | 0.75 | 0.625 |  |  |  | n/a | 180 | 0.001 | 0.05 | 0.02 | 0.02 | 20 | 1000000000000 | 19.5 | 32.5 |  |  |  | Not Tested | 1 | Not tested | [9 empty] | Not Tested | Not Tested | 26 | FALSE | Beal | 2020-04-01
67400-01 | 0.2 | 0.3 | 0.25 |  |  |  | n/a | 0 | 0.003 | 0.05 | 0.06 | 0.014 | 3 | 1000000000000 | 17 | 24 |  |  |  | 200 | 0.003 | 1 | [9 empty] | 0.935 | 1.265 | 20.5 | FALSE | Blake | 2020-03-26
67400-02 - PRELIMINARY |  |  | 0 |  |  |  | n/a |  | 0.003 | 0.05 | 0.06 | 0.014 | 3 | 1000000000000 | 17 | 24 |  |  |  | 200 | 0.003 | 1 | [11 empty] | 20.5 | FALSE | Blake | 2020-03-26
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
[11 empty] | #DIV/0! | [23 empty] | TRUE
Note on E3: Anthony J. Agnello: NOTE: If trim range is the same for the entire process, leave intial trim range blank.
Note on G3: Anthony J. Agnello: NOTE: If trim target is the same for the entire process, leave intial trim range blank.
Note on I5: Anthony J. Agnello: In Pahse = 0° Out of Phae = 180°
Note on E7: Anthony J. Agnello: TRIM MAX?
Note on J7: Anthony J. Agnello: Actual value (0.003"=0.050mV/G) has been multiplied by 1000 to force max trim ((3"=0.050mV/G).
Note on A22: Anthony J. Agnello: Placebo element. Should have no output.
Note on AG22: Christopher Kopra: Unsure of this . . . Might be for 116B - which the 46282 references in the router
Note on C26: Christopher Kopra: Need to update Router - still in mV
Note on V30: Anthony J. Agnello: mV/G

## Engineering Analysis
START HERE |  | START HERE, 0
0
| Acceleration Compensation |  |  | Diaphragm Shift
Symbol # | By: | By: | By: | Output after last trim | Final Output
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0 | *All yellow cells must be filled in
| #N/A | #N/A | #N/A |  | 0 | Initial
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
| #N/A | #N/A | #N/A |  | 0
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.