130 Microphone Calibration System
 Verification Procedure

Purpose:

To monitor and record sound pressure levels necessary to verify the system is operating properly.  This process should be applied on a regular, daily, as used, basis.

Responsibilities:

The Product Engineer is responsible to maintain this procedure and the Test Technicians are responsible to apply the procedure.

Affected Department / Product Group / Support Group:

SVS, PCB-NC

Associated Documents:

ISO9001, ISO10012,QSM and QAM

Procedure:

- The equipment listed below is used during this test and is calibrated per ISO 10012-1 and former MIL-STD-45662A.  The calibration records are on file.

- Reference 130 microphone

- Reference microphone signal conditioner (PCB 442A101 or equivalent)

- Test microphone signal conditioner (PCB 443B102 or equivalent)

- National Instruments data acquisition card

  - The equipment listed below is used during this test and no calibration is required.

- GRAS 51AB calibrator

- Computer

- Break out box

- All necessary cabling and adapters

C.	130 Microphone calibration system verification procedure

- Open the 130 Mic Cal Station Verification (TC027) file located on the network.

- Perform a calibration per AT603-2 using the reference microphone indicated at the top of the SYSCHECK worksheet.

- Record the TECHNICIAN, DATE, and SENSITIVITY in the SYSCHECK worksheet.

  - If the sensitivity is inside the yellow limit lines, the calibration system is in spec and ready to be used.

  - If the sensitivity is outside the yellow limits and inside the red limits for 7 readings in a row, notify engineering.

  - If the sensitivity is outside the red limits, verify that the stand is configured properly per AT603-2.  Repeat the measurement.  If still outside the red limits, notify engineering.

Referenced Documents:

AT603-2, TC027

Unusual Conditions:

None